IEC 62304

Edition 1.1	2015-06

CONSOLIDATED

VERSION

colour inside

Medical device software – Software life cycle processes

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

THIS PUBLICATION IS COPYRIGHT PROTECTED

Copyright © 2015 IEC, Geneva, Switzerland

All rights reserved. Unless otherwise specified, no part of this publication may be reproduced or utilized in any form       or by any means, electronic or mechanical, including photocopying and microfilm, without permission in writing from  either IEC or IEC's member National Committee in the country of the requester. If you have any questions about IEC copyright or have an enquiry about obtaining additional rights to this publication, please contact the address below or your local IEC member National Committee for further    information.

IEC Central Office

3, rue de Varembé CH-1211 Geneva 20 Switzerland

Tel.: +41 22 919 02  11

Fax: +41 22 919 03  00

info@iec.ch www.iec.ch

About the IEC

The International Electrotechnical Commission (IEC) is the leading global organization that prepares and publishes International Standards for all electrical, electronic and related   technologies.

About IEC publications

The technical content of IEC publications is kept under constant review by the IEC. Please make sure that you have the latest edition, a corrigenda or an amendment might have been   published.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC Catalogue - webstore.iec.ch/catalogue

The stand-alone application for consulting the entire bibliographical information on IEC International Standards, Technical Specifications, Technical Reports and other documents. Available for PC, Mac OS, Android Tablets and iPad.

IEC publications search - www.iec.ch/searchpub

The advanced search enables to find IEC publications by a variety of criteria (reference number, text, technical committee,…). It also gives information on projects, replaced and withdrawn publications.

IEC Just Published - webstore.iec.ch/justpublished

Stay up to date on all new IEC publications. Just Published details all new publications released. Available online  and also once a month by email.

Electropedia - www.electropedia.org

The world's leading online dictionary of electronic and electrical terms containing more than 30 000 terms and definitions in English and French, with equivalent terms in 15 additional languages. Also known as the International Electrotechnical Vocabulary (IEV) online.

IEC Glossary - std.iec.ch/glossary

More than 60 000 electrotechnical terminology entries in English and French extracted from the Terms and Definitions clause of IEC publications issued since 2002. Some entries have been collected from earlier publications of IEC TC 37, 77, 86 and CISPR.

IEC Customer Service Centre - webstore.iec.ch/csc

If you wish to give us your feedback on this publication or need further assistance, please contact the Customer Service Centre: csc@iec.ch.

Warning! Make sure that you obtained this publication from an authorized distributor.

IEC 62304

Edition 1.1

2015-06

CONSOLIDATED

VERSION

colour inside

Medical device software – Software life cycle processes

INTERNATIONAL

ELECTROTECHNICAL COMMISSION

ICS 11.040

ISBN 978-2-8322-2765-7

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304

Edition 1.1	2015-06

REDLINE VERSION

colour inside

Medical device software – Software life cycle processes

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 2 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

## CONTENTS

FOREWORD	4

INTRODUCTION	6

INTRODUCTION to Amendment 1	8

- Scope	9

  - * Purpose	9

  - * Field of application	9

  - Relationship to other standards	9

  - Compliance	9

- * Normative references	10

- * Terms and definitions	10

- * General requirements	15

  - * Quality management system	15

  - * RISK MANAGEMENT	16

  - * Software safety classification	16

  - * LEGACY SOFTWARE	18

- Software development PROCESS	19

  - * Software development planning	19

  - * Software requirements analysis	21

  - * Software ARCHITECTURAL design	23

  - * Software detailed design	24

  - * SOFTWARE UNIT implementation and verification	25

  - * Software integration and integration testing	25

  - * SOFTWARE SYSTEM testing	27

  - * Software release	28

- Software maintenance PROCESS	29

  - * Establish software maintenance plan	29

  - * Problem and modification analysis	30

  - * Modification implementation	31

- * Software RISK MANAGEMENT PROCESS	31

  - * Analysis of software contributing to  hazardous situations	31

  - RISK CONTROL measures	32

  - VERIFICATION of RISK CONTROL measures	32

  - RISK MANAGEMENT of software changes	33

- * Software configuration management PROCESS	33

  - * Configuration identification	33

  - * Change control	33

  - * Configuration status accounting	34

- * Software problem resolution PROCESS	34

  - Prepare PROBLEM REPORTS	34

  - Investigate the problem	35

  - Advise relevant parties	35

  - Use change control process	35

  - Maintain records	35

  - Analyse problems for trends	35

  - Verify software problem resolution	35

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 3 –

+AMD1:2015 CSV  IEC  2015

  - Test documentation contents	36

Annex A (informative)  Rationale for the requirements of  this standard	37

Annex B (informative)  Guidance on the provisions of  this standard	40

Annex C (informative)  Relationship to other standards	58

Annex D (informative)  Implementation	84

Bibliography	86

Index of defined terms	88

Figure 1 – Overview of software development PROCESSES  and ACTIVITIES	7

Figure 2 – Overview of software maintenance PROCESSES   and ACTIVITIES	7

Figure 3 – Assigning software safety classification	16

Figure B.2 – Pictorial representation of the relationship of HAZARD, sequence   of

events, HAZARDOUS SITUATION, and HARM – from ISO 14971:2007  Annex E	44

Figure B.1 – Example of partitioning of  SOFTWARE ITEMS	46

Figure C.1 – Relationship of key MEDICAL DEVICE standards to   IEC 62304	59

Figure C.2 – Software as part of  the V-model	62

Figure C.3 – Application of IEC 62304 with  IEC 61010-1	72

Table A.1 – Summary of requirements by software  safety class	39

Table B.1 – Development (model) strategies as defined in  ISO/IEC 12207	41

Table C.1 – Relationship to ISO 13485:2003	60

Table C.2 – Relationship to ISO  14971:2000 2007	61

Table C.3 – Relationship to IEC 60601-1	64

Table C.4 –  Relationship to IEC 60601-4   ...............................................................................

Table C.5 – Relationship to ISO/IEC 12207	74

Table D.1 – Checklist for small companies without a  certified QMS	85

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 4 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

## INTERNATIONAL ELECTROTECHNICAL COMMISSION

MEDICAL DEVICE SOFTWARE –

SOFTWARE LIFE CYCLE PROCESSES

## FOREWORD

- The International Electrotechnical  Commission (IEC) is  a  worldwide organization for standardization comprising all national electrotechnical committees (IEC National Committees).  The  object  of IEC  is to  promote international co-operation on  all questions concerning standardization in the electrical and electronic fields. To   this end and in addition to other activities, IEC publishes International Standards, Technical Specifications, Technical Reports, Publicly Available Specifications (PAS) and Guides (hereafter referred to as “IEC Publication(s)”). Their preparation  is  entrusted to  technical  committees;  any  IEC National Committee interested in the subject dealt with may participate in this preparatory work. International, governmental and non-  governmental organizations liaising with  the  IEC also participate in  this  preparation.  IEC collaborates closely  with the International Organization for Standardization (ISO) in accordance with conditions determined by agreement between the two  organizations.

- The formal decisions or agreements of IEC on technical matters express, as nearly as possible, an international consensus of opinion on the relevant subjects since each technical committee has representation from all interested IEC National  Committees.

- IEC Publications have the form of recommendations for international use and are accepted by IEC National Committees in that sense. While all reasonable efforts are made to ensure that the technical content of IEC Publications is accurate, IEC cannot be held responsible for the way in which they are used or for any misinterpretation by any end  user.

- In order to promote international uniformity, IEC National Committees undertake to apply IEC Publications transparently to the maximum extent possible in their national  and  regional  publications.  Any  divergence  between any IEC Publication and the corresponding national or regional publication shall be clearly indicated in    the latter.

- IEC itself does not provide any attestation of conformity. Independent certification bodies provide conformity assessment services and, in some areas, access to IEC marks of conformity. IEC is not responsible for  any services carried out by independent certification   bodies.

- All users should ensure that they have the latest edition of this publication.

- No liability shall attach to IEC or its directors, employees, servants or agents including individual experts and members of its technical committees and IEC National Committees for any personal injury, property damage or other damage of any nature whatsoever, whether direct or indirect, or for costs (including  legal  fees)  and  expenses arising out of the publication, use of, or reliance upon, this IEC Publication or any  other  IEC  Publications.

- Attention is drawn to the Normative references cited in this publication. Use of the referenced publications is indispensable for the correct application of this   publication.

- Attention is drawn to the possibility that some of the elements of  this  IEC Publication may be  the  subject  of  patent rights. IEC shall not be held responsible for identifying any or all such patent rights.

DISCLAIMER

This Consolidated version is not an official IEC Standard and has been prepared for user convenience. Only the current versions of the standard and its amendment(s) are to be considered the official documents.

This Consolidated version of IEC 62304 bears the edition number 1.1. It consists of the first edition (2006-05) [documents 62A/523/FDIS and 62A/528/RVD] and its amendment 1 (2015-06) [documents 62A/1007/FDIS and 62A/1014/RVD]. The technical content is identical to the base edition and its amendment.

In this Redline version, a vertical line in the margin shows where the technical content is modified by amendment 1. Additions and deletions are displayed in red, with deletions being struck through. A separate Final version with all changes accepted is available in this publication.

International Standard IEC 62304 has been prepared by a joint working group of subcommittee 62A:  Common  aspects  of  electrical  equipment  used  in  medical  practice,  of  IEC   technical

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 5 –

+AMD1:2015 CSV  IEC 2015

committee 62: Electrical equipment in medical practice and ISO Technical Committee 210, Quality management and corresponding general aspects for MEDICAL DEVICES. Table C.5 was prepared by ISO/IEC JTC 1/SC 7, Software and system  engineering.

It is published as a dual logo  standard.

This publication has been drafted in accordance with the ISO/IEC Directives, Part 2. In this standard the following print types are  used:

- requirements and definitions: in roman type;

- informative material appearing outside of tables, such as notes, examples and references:  in smaller type. Normative text of tables is also in a smaller  type;

- terms used throughout this standard that have been defined in Clause 3 and also given in   the index: in small capitals.

An asterisk (*) as the first character of a title or at the beginning of a paragraph indicates that there is guidance related to that item in Annex  B.

The committee has decided that the contents of the base publication and its amendment will remain unchanged until the stability date indicated on the IEC web site under "http://webstore.iec.ch" in the data related to the specific publication. At this date, the publication will be

- reconfirmed,

- withdrawn,

- replaced by a revised edition, or

- amended.

NOTE The attention of National Committees is drawn to the fact that equipment MANUFACTURERS and testing organizations  may  need  a   transitional  period  following  publication  of   a   new,   amended  or   revised  IEC  or   ISO publication in which to make products in accordance with the new requirements and to equip themselves for conducting new or revised tests. It is the recommendation of the committee that the content of this publication be adopted for mandatory implementation nationally not earlier than 3 years from the date of publication.

IMPORTANT – The 'colour inside' logo on the cover page of  this publication indicates that it contains colours which are considered to be useful for the correct understanding of its contents. Users should therefore print this document using a colour  printer.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 6 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

## INTRODUCTION

Software is often an integral part of MEDICAL DEVICE technology. Establishing the SAFETY and effectiveness of a MEDICAL DEVICE  containing software requires knowledge of what the software  is intended to do and demonstration that the use of the software fulfils those intentions without causing any unacceptable RISKS.

This standard provides a framework of life cycle PROCESSES with ACTIVITIES and  TASKS  necessary for the safe design and maintenance of MEDICAL DEVICE SOFTWARE. This standard provides requirements for each life cycle PROCESS. Each life cycle PROCESS  is further divided  into consists of a set of ACTIVITIES, with most ACTIVITIES further divided into consisting of a set   of TASKS.

As a basic foundation it is assumed that MEDICAL DEVICE SOFTWARE  is  developed  and  maintained within a quality management system (see 4.1) and a RISK MANAGEMENT system (see 4.2). The RISK MANAGEMENT PROCESS is already very well addressed by the  International Standard ISO 14971. Therefore IEC 62304 makes use of this advantage simply by a normative reference to ISO 14971. Some minor additional RISK MANAGEMENT requirements are needed for software, especially in the area of identification of contributing software factors related to HAZARDS. These requirements are summarized and captured in Clause 7 as the software RISK MANAGEMENT PROCESS.

Whether software is a contributing factor to a HAZARD  HAZARDOUS  SITUATION  is  determined during the HAZARD identification ACTIVITY of the RISK MANAGEMENT PROCESS. HAZARDS  HAZARDOUS SITUATIONS that could be indirectly caused by software (for example, by providing misleading information that could cause inappropriate treatment to be administered) need to be considered when determining whether software is a contributing factor. The decision to use software to control RISK is made during the RISK CONTROL ACTIVITY of the RISK MANAGEMENT PROCESS. The software RISK MANAGEMENT PROCESS required in this standard has to  be embedded in the device RISK MANAGEMENT PROCESS according to ISO   14971.

The software development PROCESS consists of a number of ACTIVITIES. These ACTIVITIES are shown in Figure 1 and described in Clause 5. Because many incidents in the field are related to service or maintenance of MEDICAL DEVICE SYSTEMS  including inappropriate software updates  and upgrades, the software maintenance PROCESS is considered to be as important as the software development PROCESS. The software maintenance PROCESS is very similar to the software development PROCESS. It is shown in Figure 2 and described in Clause   6.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 7 –

+AMD1:2015 CSV  IEC 2015

SYSTEM development ACTIVITIES (including RISK  MANAGEMENT)

5.1

Software development planning

5.2

Software requirements analysis

5.3

Software

ARCHITECTURAL

design

5.4

Software detailed design

5.5

Software UNIT implementation and VERIFICATION

5.6

Software integration and integration testing

5.7

Software SYSTEM

testing

5.8

Software release

Activities outside the scope of this standard

Customer needs

Customer needs satisfied

7  Software RISK MANAGEMENT

8  Software configuration management

9  Software problem resolution

IEC  722/06

Figure 1 – Overview of software development PROCESSES and  ACTIVITIES

Maintenance request

System maintenance ACTIVITIES (including RISK MANAGEMENT)

6.1

Establish software maintenance

plan

6.2

Problem and modification analysis

5.3

Software

ARCHITECTURAL

design

5.4

Software detailed design

5.5

Software UNIT

implementation and

VERIFICATION

5.6

Software integration and integration testing

5.7

Software SYSTEM

testing

5.8

Software release

Activities outside the scope of this standard

Request satisfied

7 Software RISK MANAGEMENT

6.3 Modification implementation

8 Software configuration management

9 Software problem resolution

IEC  723/06

Figure 2 – Overview of software maintenance PROCESSES and  ACTIVITIES

This  standard  identifies  two  additional  PROCESSES  considered  essential  for  developing safe

MEDICAL DEVICE SOFTWARE. They are the software configuration management PROCESS     (Clause

8) and the software problem resolution PROCESS (Clause  9).

Amendment 1 updates the standard to add requirements to deal with LEGACY SOFTWARE, where the software design is prior to the existence of the current version, to assist manufacturers who must  show  compliance  to   the   standard  to   meet  European  Directives.  Software      safety

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 8 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

classification changes include clarification of requirements and updating of the software safety classification to include a risk-based approach.

This standard does not specify an organizational structure for the MANUFACTURER  or which part   of the organization is to perform which PROCESS, ACTIVITY, or TASK. This standard requires only that the PROCESS, ACTIVITY, or TASK be completed to establish compliance with this   standard.

This standard does not prescribe the name, format, or explicit content of the documentation to   be produced. This standard requires documentation of TASKS, but the decision of how  to  package this documentation is left to the user of the  standard.

This standard does not prescribe a specific life cycle model. The users of this standard are responsible for selecting a life cycle model for the software project and for mapping the PROCESSES, ACTIVITIES, and TASKS in this standard onto that   model.

Annex A provides rationale for the clauses of this standard. Annex B provides guidance on the provisions of this standard.

For the purposes of this standard:

- “shall” means that compliance with a requirement is mandatory for compliance with this standard;

- “should” means that compliance with a requirement is recommended but is not mandatory for compliance with this standard;

- “may” is used to describe a permissible way to achieve compliance with a  requirement;

- “establish” means to define, document, and implement; and

- where this standard uses the term “as appropriate” in conjunction with a required PROCESS, ACTIVITY, TASK or output, the intention is that the MANUFACTURER shall use the PROCESS, ACTIVITY, TASK or output unless the MANUFACTURER can document a justification for not so doing.

## INTRODUCTION to Amendment 1

The first edition of IEC 62304 was published in 2006. This amendment is intended to add requirements to deal with LEGACY SOFTWARE, where the software design is  prior  to  the existence of the current version, to assist manufacturers who must show compliance to the standard to meet European Directives. Software safety classification changes needed for this amendment include clarification of requirements and updating of the software safety classification to include a risk-based approach. Work is continuing in parallel to develop the second edition of IEC 62304.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 9 –

+AMD1:2015 CSV  IEC 2015

# MEDICAL DEVICE SOFTWARE – SOFTWARE LIFE CYCLE PROCESSES

### Scope

#### * Purpose

This standard defines the life cycle requirements for MEDICAL DEVICE SOFTWARE. The set of PROCESSES, ACTIVITIES, and TASKS described in this standard establishes a common framework for MEDICAL DEVICE SOFTWARE life cycle  PROCESSES.

#### * Field of application

This standard applies to the development and maintenance of MEDICAL DEVICE   SOFTWARE.

This standard applies to the development and maintenance of MEDICAL DEVICE SOFTWARE when software is itself a MEDICAL DEVICE or when software is an embedded or integral part of the final MEDICAL DEVICE.

NOTE 1 This standard can be used in the development  and  maintenance  of  software  that  is  itself  a  medical device.   However,  additional development activities are needed at the system level before this type of software can    be placed into service.  These system activities are not covered by this standard, but can be found in IEC 82304-11  [22].

This standard describes PROCESSES that are intended to be applied to software which executes on a processor or which is executed by other software (for example an interpreter)  which executes on a processor.

This standard applies regardless of the persistent storage device(s) used to store the software (for example: hard disk, optical disk, permanent or flash  memory).

This standard applies regardless of the method of delivery of the software (for example: transmission by network or email, optical disk, flash memory or EEPROM). The method of software delivery itself is not considered MEDICAL DEVICE  SOFTWARE.

This standard does not cover validation and final release of the MEDICAL DEVICE, even when the

MEDICAL DEVICE consists entirely of  software.

NOTE 2 If a medical device incorporates embedded software intended to be executed on a processor,  the  requirements of this standard apply to the software, including the requirements concerning software of unknown provenance (see 8.1.2).

NOTE 3 Validation and other development activities are needed at the  system  level  before  the  software  and  medical device can be placed into service.   These system activities are not covered by this standard, but can be   found in related product standards (e.g., IEC 60601-1, IEC 82304-1,    etc.).

#### Relationship to other standards

This MEDICAL DEVICE SOFTWARE life cycle standard is to be used together with other appropriate standards when developing a MEDICAL DEVICE. Annex C shows the relationship between this standard and other relevant standards.

#### Compliance

Compliance with this standard is defined as implementing all of the PROCESSES, ACTIVITIES,   and

TASKS identified in this standard in accordance with the software safety  class.

1    In preparation.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 10 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

NOTE The software safety classes assigned to each requirement are identified in the normative text following the requirement.

Compliance is determined by inspection of all documentation required by  this  standard  including the RISK MANAGEMENT FILE, and assessment of the PROCESSES, ACTIVITIES and TASKS required for the software safety class. See Annex D.

NOTE 1    This assessment could be carried out by internal or external    audit.

NOTE 2    Although the specified PROCESSES, ACTIVITIES, and TASKS  are performed, flexibility exists in the methods     of implementing these PROCESSES and performing these ACTIVITIES and   TASKS.

NOTE 3 Where any requirements contain “as appropriate” and  were not  performed,  documentation  for  the justification is necessary for this  assessment.

NOTE 4 The term “conformance” is used in ISO/IEC 12207 where the term “compliance” is used in this standard. NOTE 5    For compliance of LEGACY SOFTWARE see  4.4.

### * Normative references

The following referenced documents are indispensable for the application of this document. For dated references, only the edition cited applies. For  undated references, the latest edition of   the referenced document (including any amendments) applies.

ISO 14971, Medical devices – Application of risk management to medical  devices.

### * Terms and definitions

For the purposes of this document, the following terms and definitions  apply.

#### 3.1

ACTIVITY

a set of one or more interrelated or interacting  TASKS

#### 3.2

ANOMALY

any condition that deviates from the expected based on requirements specifications, design documents, standards, etc. or from someone’s perceptions or experiences. ANOMALIES may be found during, but not limited to, the review, test, analysis,  compilation,  or  use  of  MEDICAL DEVICE SOFTWARE PRODUCTS  or applicable documentation

NOTE   Based on [IEEE 1044:1993, definition 3.1].

#### 3.3

ARCHITECTURE

organizational structure of a SYSTEM or component [IEEE 610.12:1990]

#### 3.4

CHANGE REQUEST

a documented specification of a change to be made to a   MEDICAL DEVICE SOFTWARE PRODUCT

#### 3.5

CONFIGURATION ITEM

entity that can be uniquely identified at a given reference  point

NOTE    Based on ISO/IEC 12207:1995 2008, 3.6  4,7.

#### 3.6

DELIVERABLE

required result or output (includes documentation) of an ACTIVITY or  TASK

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 11 –

+AMD1:2015 CSV  IEC 2015

#### 3.7

EVALUATION

a systematic determination of the extent to which an entity meets its specified criteria [ISO/IEC 12207:1995 2008, 3.9 4.12]

#### 3.8

HARM

physical injury, damage, or both to the health of people or damage to property or the  environment

[ISO/IEC Guide 51:1999, definition 3.3 ISO 14971:2007,  2.2]

#### 3.9

HAZARD

potential source of HARM

[ISO/IEC Guide 51:1999, definition 3.5 ISO 14971:2007,  2.3]

#### 3.10

MANUFACTURER

natural or legal person with responsibility for designing, manufacturing, packaging, or labelling    a MEDICAL DEVICE; assembling a SYSTEM; or adapting a MEDICAL  DEVICE  before it is  placed on  the market and/or put into service, regardless  of  whether  these operations are carried out by that person or by a third party on that person’s  behalf

NOTE 1 Attention is drawn to the fact that the provisions of national  or  regional  regulations  can  apply  to  the definition of manufacturer.

NOTE 2    For a definition of labelling, see ISO 13485:2003, definition   3.6.

[ISO 14971:2000 2007, 2.6 2,8]

#### 3.11

MEDICAL DEVICE

any instrument, apparatus, implement, machine, appliance, implant, in vitro reagent or calibrator, software, material or other similar or related article, intended by the MANUFACTURER    to be used, alone or in combination, for human beings for one or more  of  the  specific purpose(s) of

- diagnosis, prevention, monitoring, treatment or alleviation of  disease,

- diagnosis, monitoring, treatment, alleviation of or compensation for an  injury,

- investigation, replacement, modification, or  support of  the anatomy or  of  a     physiological

PROCESS,

- supporting or sustaining life,

- control of conception,

- disinfection of MEDICAL DEVICES,

- providing information for medical purposes by means of in vitro examination of specimens derived from the human body,

and which does not achieve its primary intended action in or on the human body by pharmacological, immunological or metabolic means, but which may be assisted in its function by such means

NOTE 1 This definition has been developed by the Global Harmonization Task Force (GHTF). See bibliographic reference [15] (in ISO  13485:2003).

[ISO 13485:2003, definition 3.7]

NOTE 2        Some differences can occur in the definitions used in regulations of each country.

NOTE  3	In  conjunction  with  IEC 60601-1:2005  and  IEC 60601-1:2005/AMD1:2012  the  term     “medical  device” assumes the same meaning as ME EQUIPMENT or ME SYSTEM (which are defined terms of IEC 60601-1).

#### 3.12

MEDICAL DEVICE SOFTWARE

– 12 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

SOFTWARE  SYSTEM  that  has  been  developed  for  the  purpose  of  being  incorporated  into the

MEDICAL DEVICE being developed or that is intended for use as a MEDICAL DEVICE in its own right

NOTE    This includes a MEDICAL DEVICE  software product, which then is a MEDICAL DEVICE  in its own   right.

#### 3.13

PROBLEM REPORT

a record of actual or potential behaviour of a MEDICAL DEVICE SOFTWARE PRODUCT that a user or other interested person believes to be unsafe, inappropriate for the intended use or contrary to specification

NOTE 1 This standard does not require that every PROBLEM REPORT results in a change to the  MEDICAL  DEVICE SOFTWARE PRODUCT. A MANUFACTURER can reject a PROBLEM REPORT as a misunderstanding, error or insignificant  event.

NOTE 2 A PROBLEM REPORT can relate to a released MEDICAL DEVICE SOFTWARE PRODUCT or to a MEDICAL DEVICE SOFTWARE PRODUCT that is still under  development.

NOTE 3  This standard requires the MANUFACTURER  to perform extra decision making steps (see Clause 6) for a

PROBLEM REPORT relating to a released product to ensure that regulatory actions are identified and implemented.

#### 3.14

PROCESS

a set of interrelated or interacting ACTIVITIES that transform inputs into outputs [ISO 9000:2000, definition 3.4.1]

NOTE    The term “ACTIVITIES” covers use of  resources.

#### 3.15

REGRESSION TESTING

the testing required to determine that a change to a SYSTEM component has not adversely affected functionality, reliability or performance and has not introduced additional  defects

[ISO/IEC 90003:2004, definition 3.11]

#### 3.16

RISK

combination of the probability of occurrence of HARM and the severity of that   HARM

[ISO/IEC Guide 51:1999 definition 3.2 ISO 14971:2007, 2.16]

#### 3.17

RISK ANALYSIS

systematic use of available information to identify HAZARDS and to estimate the   RISK

[ISO/IEC Guide 51:1999 definition 3.10 ISO 14971:2007, 2.17]

#### 3.18

RISK CONTROL

PROCESS in which decisions are made and RISKS are reduced to, or maintained within, specified levels

[ISO 14971:2000 2007, 2.16, modified 2.19]

#### 3.19

RISK MANAGEMENT

systematic application of management policies, procedures, and practices to the TASKS of analyzing, evaluating, and controlling RISK

[ISO 14971:2000 2007, 2.18  2.22, modified – The phrase "and monitoring" has been  removed]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 13 –

+AMD1:2015 CSV  IEC 2015

#### 3.20

RISK MANAGEMENT FILE

set of records and other documents, not necessarily contiguous, that are produced by a RISK MANAGEMENT PROCESS

[ISO 14971:2000 2007, 2.19 2.23]

#### 3.21

SAFETY

freedom from unacceptable RISK

[ISO/IEC Guide 51:1999 definition 3.1 ISO 14971:2007,  2.24]

#### 3.22

SECURITY

protection of information and data so that unauthorized people persons or systems cannot read or modify them and so that an authorized persons or systems are not denied access to   them

NOTE   Based on [ISO/IEC 12207:1995 2008, 3.25 4.39].

#### 3.23

SERIOUS INJURY

injury or illness that directly or indirectly:

- is life threatening,

- results in permanent impairment of a body function or permanent damage to a body structure, or

- necessitates medical or surgical intervention to prevent permanent impairment of a body function or permanent damage to a body structure

NOTE Permanent impairment means an irreversible impairment or damage  to  a  body  structure  or  function  excluding trivial impairment or  damage.

#### 3.24

SOFTWARE DEVELOPMENT LIFE CYCLE  MODEL

conceptual structure spanning the life of the software from definition of its requirements to its release for manufacturing, which:

- identifies the PROCESS, ACTIVITIES and TASKS involved in development of a MEDICAL DEVICE SOFTWARE PRODUCT,

- describes the sequence of and dependency between ACTIVITIES and TASKS,  and

- identifies the milestones at which the completeness of specified DELIVERABLES is   verified.

NOTE    Based on ISO/IEC 12207:1995, definition  3.11

#### 3.25

SOFTWARE ITEM

any identifiable part of a computer program, i.e., source code,  object  code,  control  code, control data, or a collection of these  items

NOTE   Three terms  identify the  software decomposition. The top level is  the SOFTWARE SYSTEM. The lowest level   that is not further decomposed is the SOFTWARE UNIT. All levels of composition, including the top and bottom levels,   can be called SOFTWARE ITEMS. A SOFTWARE SYSTEM, then, is composed of one or more SOFTWARE ITEMS, and each SOFTWARE ITEM  is  composed of  one or more SOFTWARE UNITS or decomposable SOFTWARE ITEMS. The responsibility   is left to the MANUFACTURER to provide the definition and granularity of the SOFTWARE ITEMS and SOFTWARE UNITS.

NOTE 2   Based on [ISO/IEC 90003:2004, 3.14, modified and ISO/IEC 12207:2008,   4.41]

#### 3.26

SOFTWARE PRODUCT

– 14 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

set of computer programs, procedures, and possibly associated documentation and  data

[ISO/IEC 12207:1995 definition 3.26] Not used

#### 3.27

SOFTWARE SYSTEM

integrated collection of SOFTWARE ITEMS organized to accomplish a specific function or set of functions

#### 3.28

SOFTWARE UNIT

SOFTWARE ITEM that is not subdivided into other  items

NOTE SOFTWARE UNITS can be used for the purpose of software configuration  management  or  testing.  The  granularity of SOFTWARE UNITS  is defined by the MANUFACTURER (see   B.3).

#### 3.29

SOUP

#### software of unknown provenance (acronym)

SOFTWARE ITEM that is already developed and generally available and that has not been developed for the purpose of being incorporated into the MEDICAL DEVICE (also known as “off- the-shelf software”) or SOFTWARE ITEM previously developed for which adequate records of the development PROCESSES are not  available

NOTE    A MEDICAL DEVICE SOFTWARE SYSTEM in itself cannot be claimed to be   SOUP.

#### 3.30

SYSTEM

integrated composite consisting of one or more of the PROCESSES, hardware,  software,  facilities, and people, that provides a capability to satisfy a stated need or  objective

NOTE   Based on ISO/IEC [ISO/IEC 12207:1995 2008, 3.31  4.48].

#### 3.31

TASK

a single piece of work that needs to be  done

#### 3.32

TRACEABILITY

degree to which a relationship can be established between two or more products of the development PROCESS

[IEEE 610.12:1990]

NOTE   Requirements, architecture, risk  control measures, etc.  are  examples  of  deliverables  of  the    development

PROCESS.

#### 3.33

VERIFICATION

confirmation through provision of objective evidence that specified requirements have been fulfilled

NOTE 1    “Verified” is used to designate the corresponding   status.

[ISO 9000:2000, definition 3.8.4]

NOTE 2    In  design  and  development,  VERIFICATION  concerns  the  PROCESS  of  examining  the  result  of  a      given

ACTIVITY to determine conformity with the stated requirement for that   ACTIVITY.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 15 –

+AMD1:2015 CSV  IEC 2015

#### 3.34

VERSION

identified instance of a CONFIGURATION  ITEM

NOTE 1 Modification to a VERSION of a MEDICAL DEVICE SOFTWARE PRODUCT, resulting in a new VERSION, requires software configuration management  action.

NOTE 2    Based on ISO/IEC 12207:1995 2008, 3.37  4.56.

#### 3.35

HAZARDOUS SITUATION

circumstance  in  which  people,  property  or  the  environment  are  exposed  to  one  or    more

HAZARD(S)

[SOURCE: ISO 14971:2007, 2.4]

#### 3.36

LEGACY SOFTWARE

MEDICAL DEVICE SOFTWARE which was  legally placed on the market and is  still marketed today  but for which there is insufficient objective evidence that it was  developed in  compliance with  the current version of this standard

#### 3.37

RELEASE

particular VERSION of a CONFIGURATION ITEM that is made available for a specific   purpose

NOTE    Based on ISO/IEC 12207:2008, definition  4.35.

#### 3.38

RESIDUAL RISK

RISK remaining after RISK CONTROL measures have been  taken

NOTE 1    Adapted from ISO/IEC Guide 51:1999, definition  3.9.

NOTE 2 ISO/IEC Guide 51:1999, definition 3.9 uses the term “protective measures” rather than “RISK CONTROL measures.” However, in the context of this International Standard, “protective measures” are only one option for controlling RISK  as described in 6.2 [of ISO  14971:2007].

[SOURCE: ISO 14971:2007, 2.15].

#### 3.39

RISK ESTIMATION

PROCESS  used to assign values to the probability of occurrence of HARM and the severity of    that

HARM

[SOURCE: ISO 14971:2007 2.20]

#### 3.40

RISK EVALUATION

PROCESS of comparing the estimated RISK against given RISK criteria to determine the acceptability of the RISK

[SOURCE: ISO 14971:2007 2.21]

### * General requirements

#### * Quality management system

The MANUFACTURER of MEDICAL DEVICE SOFTWARE shall demonstrate the ability  to  provide MEDICAL DEVICE SOFTWARE that consistently meets customer requirements and applicable regulatory requirements.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 16 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

NOTE 1       Demonstration of this ability can be by the use of a quality management system that complies with:

- ISO 13485 [8];  or

- a national quality management system standard;   or

- a quality management system required by national   regulation.

NOTE 2 Guidance for applying quality management system requirements to  software  can  be  found  in  ISO/IEC 90003 [15].

  - * RISK MANAGEMENT

The MANUFACTURER shall apply a RISK MANAGEMENT PROCESS complying with ISO   14971.

#### * Software safety classification

- The MANUFACTURER shall assign to each SOFTWARE SYSTEM a software safety class (A, B,    or

C) according to the possible effects on RISK of HARM  to  the  patient,  operator,  or  other people resulting from a HAZARD HAZARDOUS SITUATION to which the SOFTWARE SYSTEM can contribute in a worst-case-scenario as indicated in Figure  3.

IEC

#### Figure 3 – Assigning software safety classification

The software safety classes shall initially be assigned based on severity as follows: Class A:   No injury or damage to health is possible

Class B:   Non-SERIOUS INJURY is possible

Class C:  Death or SERIOUS INJURY is  possible

If the HAZARD could arise from a failure of the SOFTWARE SYSTEM to behave as specified, the probability of such failure shall be assumed to be 100  percent.

If the RISK of death or SERIOUS INJURY arising from a  software  failure  is  subsequently reduced to an acceptable level (as defined by ISO 14971) by a hardware RISK CONTROL measure, either by reducing the consequences of the failure or by reducing the probability     of death or SERIOUS INJURY  arising from  that failure, the software safety classification may  be reduced from C to B; and if the RISK of non-SERIOUS  INJURY  arising from  a  software  failure is similarly reduced to an acceptable level by a hardware RISK CONTROL measure, the software safety classification may be reduced from B to  A.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 17 –

+AMD1:2015 CSV  IEC 2015

The SOFTWARE SYSTEM is software safety class A  if:

- the SOFTWARE SYSTEM cannot contribute to a HAZARDOUS SITUATION;   or

- the SOFTWARE SYSTEM can contribute to a HAZARDOUS SITUATION which does not result in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM.

The SOFTWARE SYSTEM is software safety class B  if:

- the SOFTWARE SYSTEM can contribute to a HAZARDOUS SITUATION which  results  in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM and the resulting possible HARM is non-SERIOUS   INJURY.

The SOFTWARE SYSTEM is software safety class C  if:

- the SOFTWARE SYSTEM can contribute to a HAZARDOUS SITUATION which  results  in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM and the resulting possible HARM is death or SERIOUS   INJURY.

For a SOFTWARE SYSTEM initially classified as software safety class B or C, the MANUFACTURER may implement additional RISK CONTROL measures external to the SOFTWARE SYSTEM (including revising the system architecture containing the SOFTWARE SYSTEM) and subsequently assign a new software safety classification to the SOFTWARE  SYSTEM.

NOTE 1 External RISK CONTROL measures can be hardware, an independent SOFTWARE SYSTEM, health care  procedures, or other means to minimize that software can contribute to a HAZARDOUS SITUATION.

NOTE 2        See ISO 14971:2007 subclause 3.2, Management Responsibilities, for the definition of risk acceptability.

- The MANUFACTURER shall assign to each SOFTWARE SYSTEM that contributes to the implementation of a RISK CONTROL measure a software safety class based on the possible effects of the HAZARD that the RISK CONTROL measure is   controlling.

Not used.

- The MANUFACTURER shall document the software safety class assigned to each SOFTWARE SYSTEM in the RISK MANAGEMENT  FILE.

- When a SOFTWARE SYSTEM is  decomposed into SOFTWARE  ITEMS, and when a  SOFTWARE  ITEM is decomposed into further SOFTWARE ITEMS, such SOFTWARE ITEMS shall inherit the software safety classification of the original SOFTWARE ITEM (or SOFTWARE  SYSTEM) unless  the MANUFACTURER documents a rationale for classification into a different software safety class (software safety classes assigned according to 4.3 a) replacing “SOFTWARE SYSTEM” with “SOFTWARE ITEM”). Such a rationale shall explain how the new SOFTWARE ITEMS are segregated so that they may be classified separately.

- The MANUFACTURER shall document the software safety class of each SOFTWARE ITEM if that class is different from the class of the SOFTWARE ITEM from which it was created by decomposition.

- For compliance with this standard, wherever a PROCESS  is required for SOFTWARE ITEMS  of  a specific classification and the PROCESS is necessarily applied when applying this standard  to a group of SOFTWARE ITEMS, the MANUFACTURER shall  use  the  PROCESSES  and  TASKS which are required by the classification of the highest-classified  SOFTWARE  ITEM  in  the group unless the MANUFACTURER documents in the RISK MANAGEMENT FILE  a  rationale for using a lower classification.

- For each SOFTWARE SYSTEM, until a software safety class is assigned,  Class  C requirements shall apply.

NOTE   In the requirements clauses and subclauses that follow, the software safety classes that the requirement   must be performed for which a specific requirement applies  are  identified following  the  requirement  in  the  form [Class . .  .].

  - * LEGACY SOFTWARE

– 18 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

#### General

As an alternative to applying Clauses 5 through 9 of this standard, compliance of LEGACY SOFTWARE may be demonstrated as indicated in 4.4.2 to  4.4.5.

    - RISK MANAGEMENT ACTIVITIES

In accordance with 4.2 of this standard, the MANUFACTURER   shall:

- assess any feedback, including post-production information, on LEGACY SOFTWARE regarding incidents and / or near incidents, both from inside its own organization and / or from users;

- perform RISK MANAGEMENT ACTIVITIES associated with continued use of the  LEGACY  SOFTWARE, considering the following aspects:

  - integration of the LEGACY SOFTWARE in the overall MEDICAL DEVICE   architecture;

  - continuing validity of RISK CONTROL measures, implemented as part of the LEGACY SOFTWARE;

  - identification of HAZARDOUS SITUATIONS associated with the continued use of the LEGACY SOFTWARE;

  - identification of potential causes of the LEGACY SOFTWARE contributing to a HAZARDOUS SITUATION;

  - definition of  RISK  CONTROL  measures for each potential cause of  the LEGACY   SOFTWARE

contributing to a HAZARDOUS  SITUATION.

#### Gap analysis

Based on the software safety class of the LEGACY SOFTWARE (see 4.3), the MANUFACTURER shall perform a gap analysis of available DELIVERABLES against those required according to 5.2, 5.3, 5.7, and Clause 7.

- The MANUFACTURER shall assess the continuing validity of available  DELIVERABLES.

- Where  gaps  are  identified,  the  MANUFACTURER  shall  EVALUATE  the  potential  reduction in

RISK resulting from the generation of the missing DELIVERABLES and associated   ACTIVITIES.

- Based on this evaluation, the MANUFACTURER shall determine the DELIVERABLES  to  be  created and associated ACTIVITIES to be performed. The minimum DELIVERABLE shall be SOFTWARE SYSTEM test records (see  5.7.5).

NOTE Such gap analysis should assure that RISK CONTROL measures, implemented  in  LEGACY  SOFTWARE,  are included in the software  requirements.

#### Gap closure activities

- The MANUFACTURER shall establish and execute a plan to generate the identified DELIVERABLES. Where available, objective evidence may be used to generate required DELIVERABLES without performing ACTIVITIES required by 5.2, 5.3, 5.7 and Clause   7.

NOTE        A plan on how to address the identified gaps can be included in a software maintenance plan (see 6.1).

- The plan shall address the use of the problem resolution PROCESS for handling problems detected in the LEGACY SOFTWARE and DELIVERABLES in accordance with Clause   9.

- Changes to the LEGACY SOFTWARE shall be performed in accordance with Clause   6.

    - Rationale for use of LEGACY  SOFTWARE

The MANUFACTURER shall document the VERSION of the LEGACY SOFTWARE together with a  rationale for the continued use of the LEGACY SOFTWARE based on the outputs of   4.4.

NOTE    Fulfilling 4.4 enables further use of LEGACY SOFTWARE in accordance with IEC   62304.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 19 –

+AMD1:2015 CSV  IEC 2015

- Software development PROCESS

#### * Software development planning

    - Software development plan

The MANUFACTURER shall establish a software development plan (or plans) for conducting the ACTIVITIES of the software development PROCESS appropriate to the scope, magnitude, and software safety classifications of the SOFTWARE SYSTEM to be developed. The sOFTWARE DEVELOPMENT LIFE CYCLE MODEL shall either be fully defined or be referenced in the plan (or plans). The plan shall address the following:

- the PROCESSES to be used in the development of the SOFTWARE SYSTEM    (see Note 4);

- the DELIVERABLES (includes documentation) of the ACTIVITIES and  TASKS;

- TRACEABILITY   between  SYSTEM   requirements,  software  requirements,  SOFTWARE  SYSTEM

test, and RISK CONTROL measures implemented in  software;

- software configuration and change management, including SOUP CONFIGURATION ITEMS and software used to support development; and

- software problem resolution for handling problems detected in the  MEDICAL  DEVICE SOFTWARE PRODUCTS, DELIVERABLES and ACTIVITIES at each stage of the life   cycle.

[Class A, B, C]

NOTE 1 The SOFTWARE DEVELOPMENT LIFE CYCLE MODEL can identify different  elements  (PROCESSES,  ACTIVITIES, TASKS and DELIVERABLES) for different SOFTWARE ITEMS according to the software safety classification of each SOFTWARE ITEM of the SOFTWARE SYSTEM.

NOTE 2  These ACTIVITIES and TASKS  can overlap or interact and can be performed iteratively or recursively. It is not  the intent to imply that a specific life cycle model should be used.

NOTE 3 Other PROCESSES are described in this standard separately from the development PROCESS. This does not imply that they must be implemented as separate ACTIVITIES and TASKS. The ACTIVITIES and TASKS of the other PROCESSES  can be integrated into the development  PROCESS.

NOTE 4  The software development plan can reference existing PROCESSES  or    define new ones.

NOTE 5     The software development plan may be integrated in an overall SYSTEM  development plan.

#### Keep software development plan updated

The MANUFACTURER  shall update the plan as development proceeds as appropriate. [Class A,   B, C]

#### Software development plan reference to SYSTEM design and  development

- As inputs for software development, SYSTEM requirements shall be referenced in the  software development plan by the  MANUFACTURER.

- In the software development plan, the MANUFACTURER shall include or reference in the software development plan procedures for coordinating the software development and the design and development validation with the system development necessary to satisfy 4.1 (such as system integration, verification, and  validation).

[Class A, B, C]

NOTE   There might not  be  a  difference between SOFTWARE SYSTEM  requirements and  SYSTEM  requirements if    the

SOFTWARE SYSTEM  is a stand alone SYSTEM  (software-only device).

#### Software development standards, methods and tools planning

The MANUFACTURER shall include or reference in the software development  plan:

- standards,

- methods, and

- tools

– 20 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

associated with the development of SOFTWARE ITEMS of class C. [Class   C]

#### Software integration and integration testing planning

The MANUFACTURER shall include or reference in the software development plan, a plan to integrate the SOFTWARE ITEMS (including SOUP) and perform testing during integration. [Class B, C]

NOTE 1          It is acceptable to combine integration testing and SOFTWARE SYSTEM  testing into a single plan and set of

ACTIVITIES.

NOTE 2    See 5.6.

    - Software VERIFICATION planning

The MANUFACTURER shall include or reference in the software development plan the following

VERIFICATION information:

- DELIVERABLES  requiring VERIFICATION;

- the required VERIFICATION TASKS for each life cycle  ACTIVITY;

- milestones at which the DELIVERABLES are VERIFIED;  and

- the acceptance criteria for VERIFICATION of the DELIVERABLES. [Class A, B, C]

    - Software RISK MANAGEMENT  planning

The MANUFACTURER shall include or reference in the software development plan, a  plan  to conduct the ACTIVITIES and TASKS of the software RISK MANAGEMENT PROCESS, including the management of RISKS relating to SOUP. [Class A, B,  C]

NOTE    See Clause 7.

#### Documentation planning

The MANUFACTURER shall include or reference in the software development plan information about the documents to be produced during the software development life cycle. For each identified document or type of document the following information shall be included or referenced:

- title, name or naming convention;

- purpose;

- intended audience of document; and

d)c)      procedures and responsibilities for development, review, approval and  modification.

[Class A, B, C]

NOTE    See Clause 8 for consideration of configuration management of   documentation.

#### Software configuration management planning

The MANUFACTURER  shall include or reference software configuration management information  in the software development plan. The software configuration management information shall include or reference:

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 21 –

+AMD1:2015 CSV  IEC 2015

- the classes, types, categories or lists of items to be  controlled;

- the software configuration management ACTIVITIES and  TASKS;

- the  organization(s)  responsible  for  performing  software  configuration  management  and

ACTIVITIES;

- their relationship with other organizations, such as software development or  maintenance;

- when the items are to be placed under configuration control;  and

- when the problem resolution PROCESS is to be used. [Class A, B, C]

NOTE    See Clause 8.

#### Supporting items to be controlled

The items to be controlled shall include tools, items or settings, used to develop the MEDICAL DEVICE SOFTWARE, which could impact the MEDICAL DEVICE SOFTWARE. [Class B,   C]

NOTE 1 Examples of such items include compiler/assembler versions, make files, batch files, and  specific  environment settings.

NOTE 2    See Clause 8.

    - Software CONFIGURATION ITEM  control  before VERIFICATION

The MANUFACTURER shall plan to place CONFIGURATION ITEMS under documented configuration management control before they are VERIFIED. [Class B,  C]

#### Identification and avoidance of common software defects

The MANUFACTURER  shall include or reference in the software development plan a  procedure  for:

- identifying categories of defects that may be introduced based on the  selected  programming technology that are relevant to their SOFTWARE SYSTEM;  and

- documenting evidence that demonstrates that these defects do not contribute to unacceptable RISK.

NOTE    See  Annex  B  of  IEC  TR  80002-1:2009  for  examples  of  categories  of  defects  or  causes  contributing to

HAZARDOUS SITUATIONS.

[Class B, C]

#### * Software requirements analysis

    - Define and document software requirements from SYSTEM  requirements

For each SOFTWARE SYSTEM of the MEDICAL DEVICE, the MANUFACTURER shall define and  document SOFTWARE SYSTEM requirements from the SYSTEM level requirements. [Class A, B,   C]

NOTE   There might not  be  a  difference between SOFTWARE SYSTEM  requirements and  SYSTEM  requirements if    the

SOFTWARE SYSTEM  is a stand alone SYSTEM  (software-only device).

#### Software requirements content

As appropriate to the MEDICAL DEVICE SOFTWARE, the MANUFACTURER shall include  in  the  software requirements:

- functional and capability requirements;

NOTE 1  Examples  include:

- performance (e.g., purpose of software, timing  requirements),

- physical characteristics (e.g., code language, platform, operating   system),

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 22 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

- computing environment (e.g., hardware, memory size, processing unit, time zone, network infrastructure) under which the software is to perform,  and

- need for compatibility with upgrades or multiple SOUP  or other device   versions.

- SOFTWARE SYSTEM inputs and  outputs;

NOTE 2  Examples  include:

- data characteristics (e.g., numerical, alpha-numeric,  format)

- ranges,

- limits, and

- defaults.

- interfaces between the SOFTWARE SYSTEM and other  SYSTEMS;

- software-driven alarms, warnings, and operator messages;

- SECURITY requirements;

NOTE 3  Examples  include:

- those related to the compromise of sensitive   information,

- authentication,

- authorization,

- audit trail, and

- communication integrity,

- system  security/malware protection.

- usability engineering requirements that are sensitive to human errors and training user interface requirements implemented by software;

NOTE 4  Examples include those related  to:

- support for manual  operations,

- human-equipment interactions,

- constraints on personnel,  and

- areas needing concentrated human  attention.

NOTE 5 Information regarding usability engineering requirements can be found in  IEC 62366-1 [21] among others  (e.g., IEC 60601-1-6 [3]).

- data definition and database requirements;

NOTE 6  Examples  include:

- form;

- fit;

- function.

- installation and acceptance requirements of the delivered MEDICAL DEVICE SOFTWARE at the operation and maintenance site or sites;

- requirements related to methods of operation and  maintenance;

- user documentation to be developed requirements related to IT-network  aspects;

NOTE 9    Examples include those related  to:

- networked alarms, warnings, and operator  messages;

- network protocols;

- handling of unavailability of network  services.

- user maintenance requirements; and

- regulatory requirements.

NOTE 10    The requirements in a) through l) can  overlap.

[Class A, B, C]

NOTE 7       All of these requirements might not be available at the beginning of the software development.

NOTE 8 Among others, ISO/IEC 9126-1 [8] ISO/IEC 25010 [12] provides information on  quality characteristics that  may be useful in defining software  requirements.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 23 –

+AMD1:2015 CSV  IEC 2015

    - Include RISK CONTROL measures in software  requirements

The MANUFACTURER shall include RISK CONTROL measures implemented in software  for  hardware failures and potential software defects in the requirements as appropriate to the MEDICAL DEVICE SOFTWARE. [Class B,  C]

NOTE  These  requirements  might  not  be  available at  the beginning of  the software development and can change  as the software is designed and RISK CONTROL  measures are further   defined.

    - Re-EVALUATE  MEDICAL DEVICE  RISK ANALYSIS

The MANUFACTURER shall re-EVALUATE the MEDICAL DEVICE RISK ANALYSIS when software requirements are established and update it as appropriate. [Class A, B,  C]

    - Update SYSTEM requirements

The MANUFACTURER shall ensure that  existing  requirements, including  SYSTEM  requirements, are re-EVALUATED and updated as appropriate as a result of the software requirements analysis ACTIVITY. [Class A, B, C]

#### Verify software requirements

The MANUFACTURER shall verify and document that the software  requirements:

- implement SYSTEM requirements including those relating to RISK   CONTROL;

- do not contradict one another;

- are expressed in terms that avoid ambiguity;

- are stated in terms that permit establishment of test criteria and performance of tests to determine whether the test criteria have been met;

- can be uniquely identified; and

- are traceable to SYSTEM requirements or other source. [Class A, B, C]

NOTE    This standard does not require the use of a formal specification   language.

  - * Software ARCHITECTURAL design

    - Transform software requirements into an ARCHITECTURE

The MANUFACTURER shall transform the requirements for the MEDICAL DEVICE SOFTWARE into a documented ARCHITECTURE that describes the software’s structure and identifies the SOFTWARE ITEMS. [Class B, C]

    - Develop an ARCHITECTURE for the interfaces of SOFTWARE   ITEMS

The MANUFACTURER shall develop and document an ARCHITECTURE  for the interfaces between  the SOFTWARE ITEMS and the components external to the SOFTWARE ITEMS (both software and hardware), and between the SOFTWARE ITEMS. [Class B,  C]

#### Specify functional and performance requirements of SOUP  item

If a SOFTWARE ITEM is identified as SOUP, the MANUFACTURER shall specify functional and performance requirements for  the  SOUP  item  that are necessary for its intended use. [Class    B, C]

    - Specify SYSTEM hardware and software required by SOUP   item

If a SOFTWARE ITEM is identified as SOUP, the MANUFACTURER shall specify the SYSTEM hardware and software necessary to support the proper operation of the SOUP item. [Class B,   C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 24 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

NOTE  Examples include processor type and speed, memory type and size, SYSTEM  software type, communication   and display software  requirements.

    - Identify segregation necessary for RISK CONTROL

The MANUFACTURER shall identify the any segregation between  SOFTWARE  ITEMS  that  is essential to necessary for RISK CONTROL, and state how to ensure that the such segregation is effective. [Class C]

NOTE   An  example of  segregation is  to have  SOFTWARE ITEMS  execute on  different processors. The effectiveness   of the segregation can be ensured by having no shared resources between the processors. Other means  of  segregation can be applied when effectiveness can be ensured by the software ARCHITECTURE design (see B.4.3).

    - Verify software ARCHITECTURE

The MANUFACTURER shall verify and document  that:

- the ARCHITECTURE of the software implements SYSTEM and software requirements including those relating to RISK  CONTROL;

- the software ARCHITECTURE is able to support interfaces between SOFTWARE ITEMS and between SOFTWARE ITEMS and hardware;  and

- the MEDICAL DEVICE ARCHITECTURE supports proper operation of any SOUP items. [Class B, C]

NOTE       A TRACEABILITY analysis of ARCHITECTURE to software requirements can be used to satisfy requirement a).

#### * Software detailed design

    - Refine SOFTWARE ARCHITECTURE Subdivide software into SOFTWARE  UNITS

The MANUFACTURER  shall refine subdvide the software ARCHITECTURE  until it  is  represented  by

SOFTWARE UNITS. [Class B, C]

NOTE    Some SOFTWARE SYSTEMS  are not divided further.

    - Develop detailed design for each SOFTWARE  UNIT

The MANUFACTURER shall develop and document a detailed design with enough detail to allow correct implementation of for each SOFTWARE UNIT of the SOFTWARE ITEM. [Class   C]

#### Develop detailed design for interfaces

The MANUFACTURER  shall develop and document a detailed design for any interfaces between  the SOFTWARE UNIT and external components (hardware or software), as well as any interfaces between SOFTWARE UNITS, detailed enough to implement each SOFTWARE UNIT and its interfaces correctly. [Class C]

#### Verify detailed design

The MANUFACTURER shall verify and document that the software detailed  design:

- implements the software ARCHITECTURE; and

- is free from contradiction with the software ARCHITECTURE. [Class C]

NOTE It is acceptable to use a TRACEABILITY analysis of ARCHITECTURE to software detailed design to satisfy requirement a).

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 25 –

+AMD1:2015 CSV  IEC 2015

  - * SOFTWARE UNIT implementation and  verification

    - Implement each SOFTWARE UNIT

The MANUFACTURER shall implement each SOFTWARE UNIT. [Class A, B,   C]

    - Establish SOFTWARE UNIT VERIFICATION  PROCESS

The MANUFACTURER shall establish strategies, methods and procedures for verifying each the SOFTWARE UNITS. Where VERIFICATION is done by testing, the test procedures  shall  be EVALUATED for correctness adequacy. [Class B, C]

NOTE        It  is  acceptable to  combine integration testing and SOFTWARE SYSTEM  testing into a  single plan and set of

ACTIVITIES.

    - SOFTWARE UNIT acceptance criteria

The MANUFACTURER shall establish acceptance criteria for SOFTWARE UNITS prior to integration into larger SOFTWARE ITEMS as appropriate, and ensure that SOFTWARE UNITS meet acceptance criteria. [Class B, C]

NOTE     Examples of acceptance criteria  are:

- does the software code implement requirements including RISK CONTROL    measures?

- is  the  software  code  free  from  contradiction  with  the  interfaces  documented  in  the  detailed  design  of  the

SOFTWARE UNIT?

- does the software code conform to programming procedures or coding    standards?

    - Additional SOFTWARE UNIT acceptance  criteria

When present in the design, the MANUFACTURER shall include additional acceptance criteria as appropriate for:

- proper event sequence;

- data and control flow;

- planned resource allocation;

- fault handling (error definition, isolation, and recovery);

- initialisation of variables;

- self-diagnostics;

- memory management and memory overflows; and

- boundary conditions. [Class C]

    - SOFTWARE UNIT VERIFICATION

The MANUFACTURER shall perform the SOFTWARE UNIT VERIFICATION and document the results. [Class B, C]

#### * Software integration and integration testing

    - Integrate SOFTWARE UNITS

The MANUFACTURER shall integrate the SOFTWARE UNITS in accordance with the integration plan (see 5.1.5). [Class B, C]

#### Verify software integration

The MANUFACTURER shall verify and record the following aspects of the software integration that the SOFTWARE  UNITS  have been integrated into SOFTWARE  ITEMS  and/or the SOFTWARE     SYSTEM

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 26 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

in accordance with the integration plan (see 5.1.5) and retain records of the evidence of such verification.

- the SOFTWARE UNITS have been integrated into SOFTWARE ITEMS and the SOFTWARE SYSTEM; and

- the hardware items, SOFTWARE ITEMS, and support for manual operations (e.g., human- equipment interface, on-line help menus, speech recognition, voice control) of the SYSTEM have been integrated into the SYSTEM.

[Class B, C]

NOTE   This VERIFICATION  is  only that the items have been integrated integration has been done according to the  plan, not that they perform as intended. This VERIFICATION is most likely implemented by some form of inspection.

#### Test integrated Software integration testing

The MANUFACTURER shall test the integrated SOFTWARE ITEMS in accordance with the integration plan (see 5.1.5) and document the results. [Class B,  C]

#### Software integration testing content

For  software  integration  testing,  the  MANUFACTURER   shall  address  whether  the    integrated

SOFTWARE ITEM performs as intended. [Class B, C]

NOTE 1  Examples to be considered  are:

- the required functionality of the  software;

- implementation of RISK CONTROL  measures;

- specified timing and other  behaviour;

- specified functioning of internal and external interfaces;  and

- testing under abnormal conditions including foreseeable   misuse.

NOTE 2      It  is  acceptable to  combine integration testing and SOFTWARE SYSTEM  testing into a  single plan and set of

ACTIVITIES.

#### Verify EVALUATE software integration test  procedures

The MANUFACTURER shall EVALUATE the integration test procedures for correctness adequacy. [Class B, C]

#### Conduct regression tests

When software items are integrated, the MANUFACTURER shall conduct REGRESSION TESTING appropriate to demonstrate that defects have not been introduced into previously integrated software. [Class B, C]

#### Integration test record contents

The MANUFACTURER shall:

- document the test result (pass/fail and a list of  ANOMALIES);

- retain sufficient records to permit the test to be repeated;  and

- identify the tester. [Class B, C]

NOTE      Requirement b) could be implemented by retaining, for  example:

- test case specifications showing required actions and expected   results;

- records of the equipment;

- records of the test environment (including software tools) used for    test.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 27 –

+AMD1:2015 CSV  IEC 2015

#### Use software problem resolution PROCESS

The MANUFACTURER shall enter ANOMALIES found during software integration and integration testing into a software problem resolution PROCESS. [Class B,  C]

NOTE    See Clause 9.

  - * SOFTWARE SYSTEM  testing

#### Establish tests for software requirements

- The MANUFACTURER shall establish and perform a set of tests, expressed as input stimuli, expected outcomes, pass/fail criteria and procedures, for conducting SOFTWARE SYSTEM testing, such that all software requirements are covered. [Class A, B,  C]

NOTE 1 It  is  acceptable to  combine integration testing and  SOFTWARE SYSTEM  testing into a  single plan and  set of ACTIVITIES. It is also acceptable to test software requirements in earlier phases.

NOTE 2 Not only separate tests for each requirement, but also tests of combinations of requirements can be performed, especially if dependencies between requirements   exist.

- The MANUFACTURER shall EVALUATE the adequacy of VERIFICATION strategies and test procedures.

#### Use software problem resolution PROCESS

The MANUFACTURER shall enter ANOMALIES found during software system testing into a software problem resolution PROCESS. [Class A, B,  C]

#### Retest after changes

When changes are made during SOFTWARE SYSTEM testing, the MANUFACTURER   shall:

- repeat tests, perform modified tests or perform additional tests, as appropriate, to verify the effectiveness of the change in correcting the problem;

- conduct testing appropriate to demonstrate that unintended side effects have not been introduced; and

- perform relevant RISK MANAGEMENT ACTIVITIES as defined in 7.4. [Class A, B, C]

    - Verify Evaluate SOFTWARE SYSTEM  testing

The MANUFACTURER shall EVALUATE the appropriateness of VERIFICATION strategies and test procedures.

The MANUFACTURER shall verify that:

a) the VERIFICATION strategies and the test procedures used are appropriate;  b)   SOFTWARE SYSTEM test procedures trace to software  requirements;

c)a)	all software requirements have been tested or otherwise  VERIFIED;

d)b) the TRACEABILITY between software requirements and tests or other VERIFICATION is  recorded; and

e)c)	test results meet the required pass/fail criteria.

[Class A, B, C]

    - SOFTWARE SYSTEM test record contents

In order to support the repeatability of tests, the MANUFACTURER shall  document:

- a reference to test case procedures showing required actions and expected  results;

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 28 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

a)b)  the test result (pass/fail and a list of  ANOMALIES);

- retain sufficient records to permit the test to be repeated; and c)   identify the tester.

- the version of software tested;

- relevant hardware and software test configurations;

- relevant test tools;

- date tested; and

- the identity of the person responsible for executing the test and recording the test   results.

[Class A, B, C]

NOTE	Requirement b) could be implemented by retaining, for   example:

- test case specifications showing required actions and expected   results;

- records of the equipment;  and

- records of the test environment (including software tools) used for    test.

  - * Software RELEASE for utilization at a SYSTEM   level

    - Ensure software VERIFICATION is complete

The MANUFACTURER shall ensure that all software VERIFICATION ACTIVITIES have been completed and the results have been EVALUATED before the software is released. [Class A, B,   C]

    - Document known residual ANOMALIES

The MANUFACTURER shall document all known residual ANOMALIES. [Class A, B,   C]

    - EVALUATE  known residual ANOMALIES

The MANUFACTURER shall ensure that all known residual ANOMALIES have been EVALUATED to ensure that they do not contribute to an unacceptable RISK. [Class B,   C]

    - Document released VERSIONS

The MANUFACTURER shall document the VERSION  of the MEDICAL DEVICE  SOFTWARE PRODUCT that  is being released. [Class A, B,  C]

#### Document how released software was created

The MANUFACTURER shall document the procedure and environment used to create the released software. [Class B, C]

#### Ensure activities and tasks are complete

The  MANUFACTURER  shall  ensure  that  all  software  development  plan  (or  maintenance plan)

ACTIVITIES and TASKS are complete along with all the associated documentation. [Class B,   C]

NOTE    See 5.1.3.b).

#### Archive software

The MANUFACTURER shall archive:

- the MEDICAL DEVICE  SOFTWARE PRODUCT and CONFIGURATION ITEMS;  and

- the documentation

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 29 –

+AMD1:2015 CSV  IEC 2015

for at least a period of time determined as the longer of: the life time of the device MEDICAL  DEVICE SOFTWARE as defined by the MANUFACTURER or a time specified by relevant regulatory requirements. [Class A, B, C]

#### Assure repeatability of software release reliable delivery of released software

The MANUFACTURER shall establish procedures to ensure that the released MEDICAL DEVICE SOFTWARE PRODUCT can be reliably delivered to the point of use without corruption or unauthorised change. These procedures shall address the production and handling of media containing the MEDICAL DEVICE  SOFTWARE PRODUCT including as appropriate:

- replication,

- media labelling,

- packaging,

- protection,

- storage, and

- delivery. [Class A, B, C]

- Software maintenance PROCESS

#### * Establish software maintenance plan

The  MANUFACTURER  shall  establish a  software maintenance plan  (or  plans)  for conducting the

ACTIVITIES and TASKS of the maintenance PROCESS.  The plan shall address the   following:

- procedures for:

  - receiving,

  - documenting,

  - evaluating,

  - resolving and

  - tracking

feedback arising after release of the MEDICAL DEVICE   SOFTWARE;

- criteria for determining whether feedback is considered to be a  problem;

- use of the software RISK MANAGEMENT  PROCESS;

- use of the software problem resolution PROCESS for analysing  and  resolving  problems arising after release of the MEDICAL DEVICE  SOFTWARE;

- use of the software configuration management PROCESS (Clause 8) for managing modifications to the existing SOFTWARE SYSTEM;  and

- procedures to EVALUATE and  implement:

  - upgrades,

  - bug fixes,

  - patches and

  - obsolescence of SOUP.

[Class A, B, C]

#### * Problem and modification analysis

    - Document and EVALUATE  feedback

#### Monitor feedback

– 30 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

The  MANUFACTURER  shall  monitor  feedback  on released MEDICAL  DEVICE  SOFTWARE  PRODUCT

from both inside its own organization and from users released for intended use. [Class A, B,   C]

      - Document and EVALUATE  feedback

Feedback shall be documented and EVALUATED to determine whether a problem exists in a released MEDICAL DEVICE SOFTWARE PRODUCT. Any such problem shall be recorded  as  a  PROBLEM REPORT (see Clause 9). PROBLEM REPORTS shall include actual or potential adverse events, and deviations from specifications. [Class A, B,  C]

      - Evaluate PROBLEM REPORT’S affects on  SAFETY

Each PROBLEM REPORT shall be EVALUATED to determine how it affects the SAFETY of a released MEDICAL DEVICE SOFTWARE PRODUCT  released for intended use (see 9.2) and whether a change  to the released that software PRODUCT is needed to address the problem. [Class A, B,   C]

#### Use software problem resolution PROCESS

The MANUFACTURER shall use the software problem resolution PROCESS (see Clause 9)  to address PROBLEM REPORTS. [Class A, B,  C]

NOTE A problem could show that a SOFTWARE SYSTEM or SOFTWARE  ITEM  has  not  been  placed  in  the  correct software safety class. The problem resolution process can suggest changes of the software safety class. When this ACTIVITY has been  done the  PROCESS  has  been  completed, any change of  safety class in the SOFTWARE SYSTEM  or its SOFTWARE ITEMS  should be made known and  documented.

    - Analyse CHANGE REQUESTS

In addition to the analysis required by Clause 9, the MANUFACTURER shall analyse each CHANGE REQUEST for its effect on the organization, released MEDICAL DEVICE  SOFTWARE  PRODUCTS  released for intended use , and SYSTEMS with which it interfaces. [Class A, B,   C]

    - CHANGE REQUEST approval

The  MANUFACTURER  shall  EVALUATE   and  approve  CHANGE  REQUESTS   which  modify  released

MEDICAL DEVICE  SOFTWARE PRODUCTS. [Class A, B, C]

    - Communicate to users and regulators

The MANUFACTURER shall identify the approved CHANGE REQUESTS that affect released MEDICAL DEVICE  SOFTWARE PRODUCTS.

As required by local regulation, the MANUFACTURER shall inform users and regulators   about:

- any problem in released MEDICAL DEVICE SOFTWARE PRODUCTS and the consequences of continued unchanged use; and

- the nature of any available changes to released MEDICAL DEVICE SOFTWARE PRODUCTS and how to obtain and install the changes.

[Class A, B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 31 –

+AMD1:2015 CSV  IEC 2015

#### * Modification implementation

    - Use established PROCESS to implement modification

The MANUFACTURER shall use the software development PROCESS (see Clause 5) or an established maintenance PROCESS to implement the modifications identify and perform any Clause 5 ACTIVITIES that need to be repeated as a result of the modification. [Class A, B, C]

NOTE     For requirements relating to RISK MANAGEMENT of software changes    see 7.4.

    - Re-release modified SOFTWARE SYSTEM

The MANUFACTURER shall release modified SOFTWARE SYSTEMS modifications according to 5.8. [Class A, B, C]

NOTE   Modifications may can be released as part of a  full re-release of a  SOFTWARE SYSTEM  or as a  modification     kit comprising changed SOFTWARE ITEMS and the necessary tools to install the  changes  as  modifications  to  an existing SOFTWARE SYSTEM.

- * Software RISK MANAGEMENT  PROCESS

#### * Analysis of software contributing to hazardous  situations

    - Identify SOFTWARE ITEMS that could contribute to a hazardous  situation

The MANUFACTURER shall identify SOFTWARE ITEMS that could contribute to a hazardous situation identified in the MEDICAL DEVICE RISK ANALYSIS ACTIVITY of ISO 14971 (see 4.2). [Class B, C]

NOTE The hazardous situation could be the direct result of software failure or the result of the failure of a  RISK CONTROL  measure that is implemented in  software.

#### Identify potential causes of contribution to a hazardous  situation

The MANUFACTURER shall identify potential causes of the SOFTWARE ITEM identified above contributing to a hazardous situation.

The MANUFACTURER shall consider potential causes including, as  appropriate:

- incorrect or incomplete specification of functionality;

- software defects in the identified SOFTWARE ITEM  functionality;

- failure or unexpected results from  SOUP;

- hardware failures or other software defects that could result in unpredictable software operation; and

- reasonably foreseeable misuse. [Class B, C]

    - EVALUATE  published SOUP  ANOMALY lists

If failure or unexpected results from SOUP is a potential cause  of the  SOFTWARE  ITEM  contributing to a hazardous situation, the MANUFACTURER shall EVALUATE as a minimum any ANOMALY  list published by the supplier of the SOUP  item  relevant to the VERSION  of the SOUP  item used in the MEDICAL DEVICE to determine if any of the known ANOMALIES  result  in  a  sequence of events that could result in a hazardous situation. [Class B,   C]

#### Document potential causes

The  MANUFACTURER   shall  document  in  the  RISK  MANAGEMENT  FILE   potential  causes  of    the

SOFTWARE ITEM contributing to a hazardous situation (see ISO 14971). [Class B,   C]

#### Document sequences of events

– 32 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

The MANUFACTURER shall document in the RISK MANAGEMENT FILE  sequences  of  events  that could result in a hazardous situation that are identified in 7.1.2. [Class B,   C]

  - RISK CONTROL measures

    - Define RISK CONTROL measures

For each potential cause of the software item contributing to a hazardous situation case documented in the RISK MANAGEMENT FILE where a SOFTWARE ITEM could contribute to a HAZARDOUS  SITUATION, the MANUFACTURER  shall define and document RISK  CONTROL  measures  in accordance with ISO 14971. [Class B,  C]

NOTE The RISK CONTROL measures can be implemented in hardware, software, the working environment or user instruction.

    - RISK CONTROL measures implemented in  software

If a RISK CONTROL measure is implemented as part of the functions of a SOFTWARE ITEM, the

MANUFACTURER shall:

- include the RISK CONTROL measure in the software  requirements;

- assign a software safety class to the to each SOFTWARE ITEM based on the possible effects    of the HAZARD that contributes to the implementation of a RISK CONTROL measure a software safety class based on the RISK that the RISK CONTROL  measure is  controlling (see 4.3 a));  and

- develop the SOFTWARE ITEM in accordance with Clause 5. [Class B, C]

NOTE        This requirement provides additional detail for RISK CONTROL requirements of ISO 14971

  - VERIFICATION of RISK CONTROL  measures

    - Verify RISK CONTROL measures

The implementation of each RISK CONTROL measure documented in 7.2 shall be VERIFIED, and this VERIFICATION shall be documented. The MANUFACTURER shall review the RISK CONTROL measure and determine if it could result in a new HAZARDOUS SITUATION. [Class B,   C]

#### Document any new sequences of events

If a RISK CONTROL measure is implemented as a SOFTWARE ITEM, the MANUFACTURER shall EVALUATE the RISK CONTROL measure to identify and document in the RISK MANAGEMENT FILE any new  sequences  of events  that   could   result   in   a   hazardous   situation.   [Class   B,   C]  Not used.

    - Document TRACEABILITY

The MANUFACTURER shall document TRACEABILITY of software HAZARDS as   appropriate:

- from the hazardous situation to the SOFTWARE  ITEM;

- from the SOFTWARE ITEM to the specific software  cause;

- from the software cause to the RISK CONTROL measure;  and

- from the RISK CONTROL measure to the VERIFICATION of the RISK CONTROL measure. [Class B, C]

NOTE    See ISO 14971 – RISK MANAGEMENT  report.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 33 –

+AMD1:2015 CSV  IEC 2015

  - RISK MANAGEMENT of software  changes

    - Analyse changes to MEDICAL DEVICE SOFTWARE with respect to  SAFETY

The MANUFACTURER shall analyse changes to the MEDICAL DEVICE SOFTWARE (including SOUP) to determine whether:

- additional potential causes are introduced contributing to a hazardous situation;  and

- additional software RISK CONTROL measures are required. [Class A, B, C]

    - Analyse impact of software changes on existing RISK CONTROL  measures

The MANUFACTURER shall analyse changes to the software, including changes to SOUP, to determine whether the software modification could interfere with existing RISK CONTROL measures. [Class B, C]

    - Perform RISK MANAGEMENT ACTIVITIES based on  analyses

The MANUFACTURER  shall perform relevant RISK MANAGEMENT  ACTIVITIES  defined in 7.1, 7.2   and

7.3 based on these analyses. [Class B,  C]

### * Software configuration management  PROCESS

#### * Configuration identification

    - Establish means to identify CONFIGURATION  ITEMS

The MANUFACTURER shall establish a scheme for the unique identification of CONFIGURATION ITEMS and their VERSIONS to be controlled for the project according to the development and configuration planning specified in 5.1. This scheme shall include other SOFTWARE PRODUCTS or entities such as SOUP and documentation. [Class A, B,  C]

    - Identify SOUP

For each SOUP  CONFIGURATION ITEM being used, including standard libraries, the  MANUFACTURER

shall document:

- the title,

- the MANUFACTURER, and

- the unique SOUP designator

of each SOUP CONFIGURATION ITEM being used. [Class A, B,   C]

NOTE The unique SOUP designator could be, for example, a  VERSION,  a  release  date,  a  patch  number  or  an upgrade designation.

#### Identify SYSTEM configuration documentation

The MANUFACTURER shall document the set of CONFIGURATION ITEMS and their VERSIONS that comprise the SOFTWARE SYSTEM configuration. [Class A, B,  C]

#### 8.2   * Change control

8.2.1   Approve CHANGE REQUESTS

The MANUFACTURER  shall change CONFIGURATION ITEMS  identified to be controlled according    to

8.1 only in response to an approved CHANGE REQUEST. [Class A, B,   C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 34 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

NOTE 1 The decision to approve a CHANGE REQUEST can be integral to  the  change  control  PROCESS  or  part  of another PROCESS. This subclause only requires that approval of a change precede its implementation.

NOTE 2 Different acceptance PROCESSES can be used for CHANGE REQUESTS at different stages of the life cycle, as stated in plans, see 5.1.1 e) d) and 6.1   e).

#### Implement changes

The MANUFACTURER shall implement the change as specified in the CHANGE REQUEST. The MANUFACTURER shall identify and perform any ACTIVITY that needs to be repeated as a result of the change, including changes to the software safety classification of SOFTWARE SYSTEMS and SOFTWARE ITEMS. [Class A, B,  C]

NOTE   This subclause states how the change should be implemented to achieve adequate change control. It does    not imply that the implementation is an integral part of the  change control PROCESS. Implementation should use  planned PROCESSES, see 5.1.1 e) and 6.1  e).

#### Verify changes

The MANUFACTURER shall verify the change, including repeating any VERIFICATION that has been invalidated by the change and taking into account 5.7.3 and 9.7. [Class A, B,   C]

NOTE	This  subclause only requires that changes be  VERIFIED. It  does  not  imply that VERIFICATION  is    an integral part of the change control PROCESS. VERIFICATION should use planned PROCESSES, see 5.1.1 e) and 6.1 e).

    - Provide means for TRACEABILITY of  change

The MANUFACTURER shall create an audit trail whereby each maintain records  of the relationships and dependencies between:

- CHANGE REQUEST;

- relevant PROBLEM REPORT; and

- approval of the CHANGE  REQUEST

can be traced. [Class A, B,  C]

#### 8.3   * Configuration status accounting

The MANUFACTURER shall retain retrievable records of the history of controlled CONFIGURATION ITEMS including SYSTEM configuration. [Class A, B,  C]

- * Software problem resolution  PROCESS

  - Prepare PROBLEM REPORTS

The MANUFACTURER shall prepare a PROBLEM REPORT for each problem detected in  a  the  MEDICAL DEVICE SOFTWARE PRODUCT. PROBLEM REPORTS shall be classified as follows: include a statement of criticality (for example, effect on performance, SAFETY, or SECURITY) as well as  other information that may aid in the resolution of the problem (for example, devices affected, supported accessories affected).

- type;

EXAMPLE 1   corrective, preventive, or adaptive to new   environment

- scope; and

EXAMPLE 2 size of change, number of device models affected, supported accessories affected,  resources involved, time to change

- criticality.

EXAMPLE 3   effect on performance, SAFETY, or  SECURITY

[Class A, B, C]

NOTE	Problems can be discovered before or after release, inside the MANUFACTURER’S organization or outside it.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 35 –

+AMD1:2015 CSV  IEC 2015

#### Investigate the problem

The MANUFACTURER shall:

- investigate the problem and if possible identify the  causes;

- EVALUATE  the problem’s relevance to SAFETY  using the software     RISK  MANAGEMENT  PROCESS

(Clause 7);

- document the outcome of the investigation and evaluation;  and

- create a CHANGE REQUEST(S) for actions needed to correct the problem, or document the rationale for taking no action.

[Class A, B, C]

NOTE A problem does not have to be corrected for the MANUFACTURER to  comply  with  the  software  problem resolution PROCESS, provided that the problem is not relevant to   SAFETY.

#### Advise relevant parties

The MANUFACTURER shall advise relevant parties of the existence of the problem,  as  appropriate.

[Class A, B, C]

NOTE   Problems can be discovered before or after release, inside the MANUFACTURER’S  organisation or outside it.    The MANUFACTURER  determines the relevant parties depending on the   situation.

#### Use change control process

The MANUFACTURER shall approve and implement all CHANGE REQUESTS, observing the requirements of the change control PROCESS (see 8.2). [Class A, B,   C]

#### Maintain records

The MANUFACTURER shall maintain records of PROBLEM REPORTS and their resolution including their VERIFICATION.

The MANUFACTURER  shall update the RISK MANAGEMENT FILE  as appropriate (see 7.4). [Class A, B, C]

#### Analyse problems for trends

The MANUFACTURER shall perform    analysis to detect trends in PROBLEM REPORTS. [Class A, B, C]

#### Verify software problem resolution

The MANUFACTURER shall verify resolutions to determine  whether:

- problem has been resolved and the PROBLEM REPORT has been  closed;

- adverse trends have been reversed;

- CHANGE REQUESTS have been implemented in the appropriate MEDICAL DEVICE SOFTWARE PRODUCTS  and ACTIVITIES; and

- additional problems have been introduced. [Class A, B, C]

#### Test documentation contents

When testing, retesting or REGRESSION TESTING SOFTWARE ITEMS and SYSTEMS  following  a change, the MANUFACTURER shall include in the test  documentation:

a)   test results;

– 36 –

IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

b)

c)

d)

e)

f)

g)

ANOMALIES  found;

the VERSION of software tested;

relevant hardware and software test configurations; relevant test tools;

date tested; and

identification of the tester.

[Class A, B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 37 –

+AMD1:2015 CSV  IEC 2015

# Annex A

## (informative)

Rationale for the requirements of this  standard

Rationale for the clauses of this standard is provided in this  annex.

### Rationale

The primary requirement of this standard is that a set of PROCESSES be followed in the development and maintenance of MEDICAL DEVICE SOFTWARE, and that the choice of PROCESSES be appropriate to the RISKS to the patient and other people. This follows from the belief that testing of software is not sufficient to determine that it is safe in   operation.

The PROCESSES required by this standard fall into two  categories:

- PROCESSES which are required to determine the RISKS  arising from the operation of each

SOFTWARE ITEM in the software;

- PROCESSES which are required to achieve an appropriately low probability of software failure for each SOFTWARE ITEM, chosen on the basis of these determined   RISKS.

This standard requires the first category to be performed for all MEDICAL DEVICE SOFTWARE  and   the second category to be performed for selected SOFTWARE  ITEMS.

A claim of compliance with this standard should therefore include a documented RISK ANALYSIS that identifies foreseeable sequences of events that include software and that can result in a HAZARDOUS  SITUATION  (see ISO 14971). HAZARDS HAZARDOUS  SITUATIONS  that can be  indirectly

caused by software (for example, by providing misleading information that could cause inappropriate treatment to be administered) should be included in this RISK   ANALYSIS.

All ACTIVITIES that are required as part of the first category of PROCESSES are identified in the normative text as "[Class A, B, C]", indicating that they are required irrespective of the classification of the software to which they  apply.

ACTIVITIES are required for all classes A, B, and C for the following   reasons:

- the ACTIVITY produces a plan relevant to RISK MANAGEMENT or software   safety classification;

- the ACTIVITY produces an output that is an input to RISK MANAGEMENT or software safety classification;

- the ACTIVITY is a part of RISK MANAGEMENT or software safety   classification;

- the ACTIVITY establishes an administration  system,  documentation  or record-keeping system that supports RISK MANAGEMENT or software safety  classification;

- the ACTIVITY normally takes place when the classification of the related  software  is unknown;

- the ACTIVITY can cause a change that could invalidate the current software safety classification of the associated software. This includes the discovery and analysis of safety related problems after release.

Other PROCESSES are required only for SOFTWARE SYSTEMS or SOFTWARE ITEMS classified in software safety classes B or C. ACTIVITIES  required as parts of these PROCESSES  are identified  in the normative text as "[Class B, C]", or "[Class C]" indicating that they  are  required selectively depending on the classification of the software to which they  apply.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 38 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

ACTIVITIES are required selectively for software in classes B and C for the following   reasons:

- the ACTIVITY enhances the reliability of the software by requiring more detail or more rigor in the design, testing or other VERIFICATION;

- the ACTIVITY is an administrative ACTIVITY that supports another  ACTIVITY  required  for  classes B or C;

- the ACTIVITY supports the correction of safety-related  problems;

- the ACTIVITY produces records of design, implementation, VERIFICATION and release of safety-related software.

ACTIVITIES are required selectively for software in class C for the following   reasons:

- the ACTIVITY further enhances the reliability of  the  software  by requiring more detail, or more rigour, or attention to specific issues in the design, testing or other   VERIFICATION

Note that all PROCESSES and ACTIVITIES defined in this standard are considered valuable in assuring the development and maintenance of high quality software. The omission of many of these PROCESSES and ACTIVITIES as requirements for software in class A that cannot  by  definition cause a  HAZARD should not imply that these PROCESSES  and ACTIVITIES would not be    of value or are not recommended. Their omission is intended to recognize that software that cannot cause a HAZARD can be easily assured of SAFETY and effectiveness primarily through overall validation ACTIVITY during the design of a MEDICAL DEVICE (which is outside the scope of this standard) and through some simple software life cycle  controls.

### Summary of requirements by class

Table A.1 summarizes which software safety classes are assigned to each requirement. This table is informative and only provided for convenience. The normative section identifies the software safety classes for each requirement.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 39 –

+AMD1:2015 CSV  IEC 2015

#### Table A.1 – Summary of requirements by software safety  class

| Clauses and subclauses |  |  |  |  | Class A | Class B | Class C |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | 4 | All requirements |  |  | X | X | X |
| 5.1 |  | 5.1.1, 5.1.2, 5.1.3, 5.1.6, 5.1.7, | 5.1.8, | 5.1.9 | X | X | X |
| 5.1.5, 5.1.10, 5.1.11, 5.1.12 |  |  |  |  |  | X | X |
| 5.1.4 |  |  |  |  |  |  | X |
| 5.2 |  | 5.2.1, 5.2.2, 5.2.4, 5.2.5, 5.2.6 |  |  | X | X | X |
| 5.2.3 |  |  |  |  |  | X | X |
| 5.3 |  | 5.3.1, 5.3.2, 5.3.3, 5.3.4, 5.3.6 |  |  |  | X | X |
| 5.3.5 |  |  |  |  |  |  | X |
| 5.4 |  | 5.4.1 |  |  |  | X | X |
| 5.4.2, 5.4.3, 5.4.4 |  |  |  |  |  |  | X |
| 5.5 |  | 5.5.1 |  |  | X | X | X |
| 5.5.2, 5.5.3, 5.5.5 |  |  |  |  |  | X | X |
| 5.5.4 |  |  |  |  |  |  | X |
| 5.6 |  | All requirements |  |  |  | X | X |
| 5.7 |  | All requirements |  |  | X | X | X |
| 5.8 |  | 5.8.1, 5.8.2, 5.8.4, 5.8.7, 5.8.8 |  |  | X | X | X |
|  |  | 5.8.1, 5.8.2, 5.8.3, 5.8.5, 5.8.6, | 5.8.7, | 5.8.8 |  | X | X |
| 6.1 |  | 6.1 |  |  | X | X | X |
| 6.2 |  | 6.2.1, 6.2.2, 6.2.4, 6.2.5 |  |  | X | X | X |
| 6.2.3 |  |  |  |  |  | X | X |
| 6.3 |  | All requirements |  |  | X | X | X |
| Clause | 6 | All requirements |  |  | X | X | X |
| 7.1 |  | All requirements |  |  |  | X | X |
| 7.2 |  | All requirements |  |  |  | X | X |
| 7.3 |  | All requirements |  |  |  | X | X |
| 7.4 |  | 7.4.1 |  |  | X | X | X |
| 7.4.2, 7.4.3 |  |  |  |  |  | X | X |
| Clause | 8 | All requirements |  |  | X | X | X |
| Clause | 9 | All requirements |  |  | X | X | X |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 40 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

# Annex B

## (informative)

Guidance on the provisions of this standard

### Scope

#### Purpose

The purpose of this standard is to provide a development PROCESS that will  consistently produce high quality, safe MEDICAL DEVICE SOFTWARE. To accomplish this, the standard  identifies the minimum ACTIVITIES and TASKS that need to be accomplished to provide confidence that the software has been developed in a manner that is likely to produce highly reliable and safe MEDICAL DEVICE SOFTWARE  PRODUCTS.

This annex provides guidance for the application of the requirements of this standard. It does not add to, or otherwise change, the requirements of this standard. This annex can be used     to better understand the requirements of this  standard.

Note that in this standard, ACTIVITIES are subclauses called out within the PROCESSES  and  TASKS are defined within the ACTIVITIES. For example, the ACTIVITIES defined for the software development PROCESS are software development planning, software requirements analysis, software ARCHITECTURAL design, software detailed design, SOFTWARE UNIT implementation and VERIFICATION, software integration and integration testing, SOFTWARE SYSTEM testing, and software release. The TASKS within these ACTIVITIES are the individual   requirements.

This standard does not require a particular SOFTWARE DEVELOPMENT LIFE CYCLE  MODEL. However, compliance with this standard does imply dependencies between PROCESSES, because inputs of a PROCESS are generated by another PROCESS. For example, the software safety classification of the SOFTWARE SYSTEM should be completed after the RISK ANALYSIS PROCESS has established what HARM could arise from failure of the   SOFTWARE SYSTEM.

Because of such logical dependencies between processes, it is easiest to describe the processes in this standard in a sequence, implying a “waterfall” or “once-through” life cycle model. However, other life cycles can also be used. Some development (model) strategies as defined at ISO/IEC 12207 [9] include (see also Table   B.1):

- Waterfall. The “once-through" strategy, also called “waterfall”, consists of performing the development PROCESS a single time. Simplistically: determine customer needs, define requirements, design the SYSTEM, implement the system, test, fix and   deliver.

- Incremental: The “incremental” strategy determines customer needs and defines the SYSTEM requirements, then performs the rest of the development in a sequence of builds. The first build incorporates part of the planned capabilities, the next build adds more capabilities, and so on, until the SYSTEM is  complete.

- Evolutionary: The “evolutionary” strategy also develops a SYSTEM in builds but differs from the incremental strategy in acknowledging that the user need is not fully understood and   all requirements cannot be defined up front. In this strategy, customer needs and SYSTEM requirements are partially defined up front, then are refined in each succeeding   build.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 41 –

+AMD1:2015 CSV  IEC  2015

#### Table B.1 – Development (model) strategies as defined in ISO/IEC   12207

| Development Strategy | Define all requirements first? | Multiple development cycles? | Distribute interim software? |
| --- | --- | --- | --- |
| Waterfall (Once-through) | yes | no | no |
| Incremental (Preplanned product improvement) | yes | yes | maybe |
| Evolutionary | no | yes | yes |

Whichever life cycle is chosen it is necessary to maintain the logical dependencies between PROCESS outputs such as specifications, design documents and software. The waterfall life cycle model achieves this by delaying the start of a PROCESS until the inputs for that PROCESS are complete and approved.

Other life cycles, particularly evolutionary life cycles, permit PROCESS outputs to be produced before all the inputs for that PROCESS are available. For example, a  new SOFTWARE  ITEM  can be specified, classified, implemented and VERIFIED before the whole software ARCHITECTURE has been finalised. Such life cycles carry the RISK that a change or  development  in  one PROCESS output will invalidate another PROCESS output. All life cycles therefore use a comprehensive configuration management system to ensure that all PROCESS outputs are brought to a consistent state and the dependencies  maintained.

The following principles are important regardless of the software development life cycle   used:

- All PROCESS outputs should be maintained in a consistent state; whenever any PROCESS output is created or changed, all related PROCESS outputs should be updated promptly to maintain their consistency with each other and to maintain all dependencies explicitly or implicitly required by this standard;

- all PROCESS outputs should be available when needed as input to further work on the software.

- before any MEDICAL DEVICE SOFTWARE is released, all PROCESS outputs  should  be consistent with each other and all dependencies between PROCESS outputs explicitly or implicitly required by this standard should be  observed.

#### Field of application

This standard applies to the development and maintenance of MEDICAL DEVICE SOFTWARE  as well as the development and maintenance of a MEDICAL DEVICE that includes   SOUP.

The use of this standard requires the MANUFACTURER to perform MEDICAL DEVICE RISK MANAGEMENT that is compliant with ISO 14971. Therefore, when the MEDICAL DEVICE SYSTEM ARCHITECTURE includes an acquired component (this could be a purchased component or a component of unknown provenance), such as a printer/plotter that  includes  SOUP,  the acquired component becomes the responsibility of the MANUFACTURER and must be included     in the RISK MANAGEMENT of the MEDICAL DEVICE. It is assumed that through proper performance   of MEDICAL DEVICE RISK MANAGEMENT, the  MANUFACTURER  would  understand  the  component and recognize that it includes SOUP. The MANUFACTURER using this standard would invoke the software RISK MANAGEMENT PROCESS as part of the overall MEDICAL DEVICE RISK MANAGEMENT PROCESS.

The maintenance of released MEDICAL DEVICE SOFTWARE applies to the post-production experience with the MEDICAL DEVICE SOFTWARE. Software maintenance includes the combination of all technical and administrative means, including supervision actions, to act     on problem reports to retain an item in, or restore it  to, a  state in which it can perform a  required  function  as   well  as   modification  requests  related  to   released  MEDICAL    DEVICE

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 42 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

SOFTWARE PRODUCT(S). For example, this includes problem rectification, regulatory reporting, re-validation and preventive action. See ISO/IEC 14764  [10].

### Normative references

ISO/IEC 90003 [15] provides guidance for applying a quality management system to software development. This guidance is not required by this standard but is highly   recommended.

### Terms and definitions

Where possible, terms have been defined using definitions from international   standards.

This standard chose to use three terms to describe the decomposition of a SOFTWARE SYSTEM (top level). The SOFTWARE SYSTEM can be a subsystem of the MEDICAL DEVICE (see IEC 60601- 1-4 [2]) or a MEDICAL DEVICE in its own right, which then becomes a software MEDICAL DEVICE.  The lowest level that is not further decomposed for the purposes of testing or software configuration management is the SOFTWARE UNIT. All levels of composition, including the  top and bottom levels, can be called SOFTWARE ITEMS. A SOFTWARE SYSTEM, then, is composed of one or more SOFTWARE ITEMS, and each SOFTWARE ITEM is composed  of  one  or  more SOFTWARE UNITS or decomposable SOFTWARE ITEMS. The responsibility is left to the MANUFACTURER to provide the definition and granularity of the SOFTWARE ITEMS and SOFTWARE UNITS. Leaving these terms vague allows one to apply them to  the  many  different  development methods and types of software used in MEDICAL   DEVICES.

### General requirements

There is no known method to guarantee 100 %    SAFETY for any kind of software.

There are three major principles which promote SAFETY for MEDICAL DEVICE    SOFTWARE:

- RISK MANAGEMENT;

- quality management;

- software engineering.

For the development and maintenance of safe MEDICAL DEVICE SOFTWARE it is necessary to establish RISK MANAGEMENT as an integral part of a quality management system as an overall framework for the application of appropriate software engineering methods  and  techniques. The combination of these three concepts allows a MEDICAL DEVICE MANUFACTURER to follow a clearly structured and consistently repeatable decision-making PROCESS  to  promote SAFETY    for MEDICAL DEVICE SOFTWARE.

#### Quality management system

A disciplined and effective set of software  PROCESSES  includes  organizational  PROCESSES such as management, infrastructure, improvement, and training. To avoid duplication and to focus this standard on software engineering, these PROCESSES have been omitted from this standard. These PROCESSES are covered by  a  quality management system. ISO  13485 [8]  is an International Standard that is specifically intended for applying the concepts of quality management to MEDICAL DEVICES. Conformance to ISO 13485 quality management system requirements does not automatically constitute conformity with national or regional regulatory requirements. It is the MANUFACTURER’S responsibility to  identify  and  establish  compliance with relevant regulatory requirements.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 43 –

+AMD1:2015 CSV  IEC  2015

    - RISK MANAGEMENT

Software development participates in RISK  MANAGEMENT  ACTIVITIES  sufficiently to ensure that all reasonably foreseeable RISKS associated with the MEDICAL DEVICE  SOFTWARE  are considered.

Rather than trying to define an appropriate RISK MANAGEMENT PROCESS in this software engineering standard, it is required that the MANUFACTURER apply  a  RISK  MANAGEMENT  PROCESS that is compliant with ISO 14971, which deals explicitly with RISK MANAGEMENT for MEDICAL DEVICES. Specific software RISK MANAGEMENT ACTIVITIES resulting from HAZARDS HAZARDOUS SITUATIONS that have software as a contributing cause are identified  in  a  supporting PROCESS described in Clause  7.

#### Software safety classification

The RISK associated with software as a part of a MEDICAL DEVICE, as  an  accessory  to  a MEDICAL DEVICE, or as a MEDICAL DEVICE in its own right, is used as the input to a  software  safety classification scheme, which then determines the PROCESSES to be used during the development and maintenance of software.

RISK is considered to be a combination of the severity of injury HARM and the probability of its occurrence. However, there is no consensus on how to determine the  probability  of  occurrence of software failures using traditional statistical methods. In this  standard, therefore, SOFTWARE SYSTEM classification is based on the severity of the  HAZARD  resulting  from failure of the software, assuming that the failure will occur. SOFTWARE SYSTEMS that contribute to the implementation of RISK CONTROL measures are classified based  on  the severity of the HAZARD they are controlling. However, no consensus exists for a method of quantitatively estimating the probability of occurrence of a software failure. When software is present in a sequence or combination of events leading to a HAZARDOUS SITUATION, the probability of the software failure occurring cannot be  considered in  estimating the  RISK  for the HAZARDOUS SITUATION. In such cases, considering a worst case probability is appropriate, and the probability for the software failure occurring should be set to 1. When it is possible to estimate the probability for the remaining events in the sequence (as it may be if they are not software) that probability can be used for the probability of  the  HAZARDOUS  SITUATION occurring (P1  in Figure B.2).

In many cases however, it might not be possible to estimate the probability for the remaining events in the sequence, and the RISK should be EVALUATED on the basis of the nature of the

HARM alone (the probability of the HAZARDOUS SITUATION occurring should be set to 1). RISK ESTIMATION in these cases should be focused on the SEVERITY of the HARM resulting from the HAZARDOUS SITUATION. Subjective rankings of probability can also be assigned based  on clinical knowledge to distinguish failures that a  clinician would be likely to detect from those that would not be detected and would be more likely to cause   HARM.

Estimates of probability of a HAZARDOUS SITUATION leading to HARM  (P2  in  Figure  B.2)  generally require clinical knowledge to distinguish between HAZARDOUS SITUATIONS where clinical practice would be likely to prevent HARM, and HAZARDOUS SITUATIONS  that would be more likely to cause HARM.

– 44 –	IEC 62304:2006

Sequence of events

+AMD1:2015 CSV  IEC 2015

NOTE	P1 is the probability of a hazardous situation   occurring

P2 is the probability of a hazardous situation leading to    harm

Hazard

Exposure (P1)

Hazardous

situation

P2

Harm

Severity of

the harm

Probability

of occurrence of harm

Risk

P1  × P2

IEC

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

Figure B.2 – Pictorial representation of the relationship of HAZARD, sequence of  events,

HAZARDOUS SITUATION, and HARM  – from ISO 14971:2007 Annex  E

If a SOFTWARE SYSTEM is decomposed into SOFTWARE ITEMS, then  each SOFTWARE  ITEM  can  have its own software safety classification.

It is only possible to determine the RISK associated with failure of a   SOFTWARE ITEM:

- if a SYSTEM ARCHITECTURE and a software ARCHITECTURE  define the role of  the SOFTWARE  ITEM in terms of its purpose and its interfaces with other software and hardware   items;

- if changes to the SYSTEM are  controlled;

- after RISK ANALYSIS has been done on the ARCHITECTURE and RISK CONTROL measures specified.

This standard requires the minimum number of ACTIVITIES that will  achieve  the  above  conditions for all classes of software.

The end of the software ARCHITECTURE ACTIVITY  is  the earliest point in the development when  the full set of SOFTWARE ITEMS is defined and the RISK MANAGEMENT ACTIVITY has identified how the SOFTWARE ITEMS relate to SAFETY. This is therefore the earliest point at  which  SOFTWARE ITEMS can be classified definitively according to their SAFETY   role.

This point corresponds to the point where RISK CONTROL is begun in ISO   14971.

Before this point, the RISK MANAGEMENT PROCESS identifies ARCHITECTURAL RISK CONTROL measures, for example adding protective subsystems, or reducing  the  opportunities  for software failures to cause HARM. After this point, the RISK MANAGEMENT  PROCESS  uses PROCESSES aimed at reducing the probability of failure of SOFTWARE ITEMS. In other words, the classification of a SOFTWARE ITEM specifies PROCESS-based RISK CONTROL measures to be  applied to that item.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 45 –

+AMD1:2015 CSV  IEC 2015

It is expected that MANUFACTURERS will find it useful to classify software before this point, for example to focus attention on areas to be investigated, but such classification should be regarded as preliminary and should not be used to justify the omission of  PROCESSES.

The software safety classification scheme is not intended to align with the RISK  classifications    of ISO 14971. Whereas the ISO 14971 scheme classifies RISK according to their severity and likelihood, the software safety classification scheme classifies SOFTWARE SYSTEMS and SOFTWARE ITEMS according to the PROCESSES to be applied in their development and maintenance.

As the design evolves, new RISKS might become evident. Therefore, RISK  MANAGEMENT  should  be applied as an integral part of the development PROCESS. This permits the development of an ARCHITECTURAL design that identifies a complete set  of  SOFTWARE  ITEMS, including those that are required to function correctly to assure safe operation and those that prevent faults from causing HARM.

The software ARCHITECTURE should promote segregation of software items that are required for safe operation and should describe the methods used to ensure effective segregation of those SOFTWARE ITEMS. Segregation is not restricted to physical (processor or memory partition) separation but includes any mechanism that prevents one SOFTWARE ITEM from negatively affecting another. The adequacy of a segregation is determined based on the RISKS  involved  and the rationale which is required to be  documented.

As stated in B.3, this standard chooses to use three terms to describe the decomposition of a

SOFTWARE SYSTEM (top level).

Figure B.1 illustrates the possible partitioning for SOFTWARE ITEMS  within a  SOFTWARE  SYSTEM and how the software safety classes would be applied to the group of SOFTWARE ITEMS in the decomposition.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 46 –	IEC 62304:2006

SOFTWARE SYSTEM /

SOFTWARE ITEM

(CLASS C)

SOFTWARE ITEM

X

(Class A)

SOFTWARE ITEM

Y

(Class C)

SOFTWARE ITEM

W

(Class B)

SOFTWARE ITEM

Z

(Class C)

IEC  724/06

+AMD1:2015 CSV  IEC 2015

Figure B.1 – Example of partitioning of SOFTWARE  ITEMS

For this example, the MANUFACTURER knows, due to  the  type  of  MEDICAL  DEVICE  SOFTWARE being developed, that the preliminary software safety classification for the SOFTWARE SYSTEM is software safety class C. During software ARCHITECTURE design the MANUFACTURER has decided  to partition the SYSTEM, as shown, with 3 SOFTWARE ITEMS – X, W and Z. The MANUFACTURER is able to segregate all SOFTWARE SYSTEM contributions to HAZARDS HAZARDOUS SITUATIONS which could result in death or SERIOUS INJURY to SOFTWARE ITEM Z and  all  remaining  SOFTWARE  SYSTEM contributions to HAZARDS HAZARDOUS SITUATIONS which could result in a non-SERIOUS INJURY to SOFTWARE ITEM W. SOFTWARE ITEm W is classified as software safety class B and SOFTWARE ITEM Z is at software safety class C. SOFTWARE ITEM  Y therefore must be classified  as Class C, per 4.3 d). The SOFTWARE SYSTEM is also at a software safety class C per this requirement. SOFTWARE ITEM X has been classified at a software safety class of A. The MANUFACTURER is able to document a rationale for the segregation between SOFTWARE ITEMS X and Y, as well as SOFTWARE ITEMS W and Z, to assure the integrity of the segregation. If partitioning segregation is not possible between SOFTWARE ITEMS  X and Y, then SOFTWARE ITEM   X must be classified in software safety class  C.

B.4.4	LEGACY SOFTWARE

Subclause 4.4 establishes a process for application of this  standard  to  LEGACY  SOFTWARE. Some geographies may require the MANUFACTURER to show conformity to the standard to obtain regulatory approval of  the MEDICAL  DEVICE SOFTWARE, even if that software was designed prior  to the existence of the current version of the standard (LEGACY SOFTWARE). In this case, the requirements in 4.4 provide a method for the the MANUFACTURER to demonstrate compliance of LEGACY SOFTWARE to the  standard.

A MANUFACTURER may determine that retrospective documentation of an already finished development-lifecycle performed as an isolated activity does not result in the reduction of RISK associated with the use of the product. The process results in the identification of a subset of

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 47 –

+AMD1:2015 CSV  IEC 2015

ACTIVITIES defined in this standard which does  result in  reduction of  RISK. Some additional  goals implicit in the process are:

- required ACTIVITIES and resulting documentation should rely on and make use of, wherever possible, existing documentation, and

- a MANUFACTURER should utilize resources as effectively as possible to effect a reduction of

RISK.

In addition to a plan identifying the subset of ACTIVITIES to execute, the process also results in objective evidence supporting safe continued use of the LEGACY SOFTWARE and a summary rationale for this conclusion.

The RISKS associated with the planned continued use of the LEGACY SOFTWARE depend on the context in which the LEGACY SOFTWARE will be used to create a SOFTWARE SYSTEM. The MANUFACTURER will document all identified MEDICAL DEVICE HAZARDS associated with the LEGACY SOFTWARE.

Subclause 4.4 requires a comprehensive assessment of available post-production field data obtained for the LEGACY SOFTWARE during the time it has been in production and use. Typical sources of post-production data include:

- adverse events attributable to the device,

- feedback received from users of the device,  and

- ANOMALIES  discovered by the MANUFACTURER.

Though no consensus exists for a method of prospectively estimating quantitatively the probability of occurrence of a software failure, such information may be available for LEGACY SOFTWARE, based on the usage of such software and EVALUATION  of post-production data. If it    is possible in such cases to quantitatively estimate the probability of events in the sequence, a quantitative value may be used for expressing the probability of the entire sequence of events occurring. If such quantitative estimation is not possible, considering a worst case probability is appropriate, and the probability for the software failure occurring should be assumed to be   1.

The MANUFACTURER determination of how the LEGACY SOFTWARE will be used in the  overall MEDICAL DEVICE SYSTEM ARCHITECTURE is input to the assessment of RISK. The RISKS to be considered vary accordingly.

- When LEGACY SOFTWARE has been safely and reliably used and the MANUFACTURER  wishes  to continue use of the LEGACY SOFTWARE, the rationale for continued use rests primarily on the assessment of RISK based on post-production  records.

- When LEGACY SOFTWARE is reused to create a new SOFTWARE SYSTEM, the intended use of the LEGACY SOFTWARE might be different from its original intended use. In this case the RISK assessment must take into account the modified set of HAZARDOUS SITUATIONS which can arise due to failures of the LEGACY  SOFTWARE.

- A reused LEGACY SOFTWARE may be used for similar intended use but integrated into a new SOFTWARE SYSTEM.  In this case the RISK  assessment should take into account modification of architectural RISK CONTROL measures according to  5.3.

When LEGACY SOFTWARE will be changed and used within a new SOFTWARE SYSTEM, the MANUFACTURER should consider how the existing records of safe and reliable operation may be invalidated by the changes.

Changes to the LEGACY SOFTWARE should be performed according to Clauses 4 to 9 of this standard, including assessment of impact to RISK CONTROL measures according to 7.4. In the case of LEGACY SOFTWARE, existing RISK CONTROL measures may not be fully documented and special care should be taken to EVALUATE the potential impact of changes, utilizing available documented design records as well as expertise of individuals having knowledge of the   system.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 48 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

According to 4.4, the MANUFACTURER performs a gap analysis in order to determine  the  available documentation including objective evidence of performed TASKS done during development of the LEGACY SOFTWARE  and compared to 5.2, 5.3, 5.7, and Clause 7. Typical  steps to accomplish this gap analysis  include

- identification of the LEGACY SOFTWARE, including VERSION, revision and any other means, required for clear identification;

- EVALUATION of existing DELIVERABLES corresponding to the deliverables required by 5.2, 5.3, 5.7, and Clause 7;

- EVALUATION of available objective evidence, documenting the previously applied software development lifecycle model (as appropriate);

- EVALUATION of the adequacy of existing RISK MANAGEMENT documentation, taking ISO 14971 into account.

Taking the performed gap analysis into account, the MANUFACTURER will EVALUATE the potential reduction in RISK resulting from the generation of the missing DELIVERABLES and associated ACTIVITIES, and create a plan to perform ACTIVITIES and generate DELIVERABLES to close these gaps.

Reduction of RISK should balance the benefit of applying the software development process according to Clause 5 against the possibility that modification of the LEGACY SOFTWARE without  full knowledge of its development history could introduce new defects that increase the  risk. Some of the elements of Clause 5 may be assessed to have little to no reduction of RISK when done after the fact. For example, detailed design and unit verification reduce RISK primarily during the process of developing new software or refactoring existing software. If these  objectives are not planned, performing the ACTIVITIES in isolation may create documentation but lead to no reduction in RISK.

At a minimum, the gap closure plan addresses missing SOFTWARE SYSTEM test records. If these  do not exist or are not suitable to support a rationale to continue use of the LEGACY SOFTWARE, the gap closure plan should include creation of SOFTWARE SYSTEM requirements at a functional level according to 5.2 and tests according to  5.7.

The documented rationale for continued use of the LEGACY SOFTWARE builds on the available objective evidence and analysis obtained in the course of  assessing the  RISK  and  creating a gap closure plan appropriate for the context of LEGACY SOFTARE   reuse.

The rationale makes a positive case for the safe and reliable performace of the LEGACY  SOFTWARE in the planned reuse context, taking into account both the post-production records available for the LEGACY SOFTWARE and the RISK CONTROL MEASURES affected by filling process gaps.

After LEGACY SOFTWARE has been re-used according to 4.4, those parts  of  the  LEGACY  SOFTWARE for which gaps in DELIVERABLES  remain, continue to be LEGACY  SOFTWARE  and may  be considered for further re-use again according to 4.4. When gaps in deliverables are closed   by changing the LEGACY  SOFTWARE, the changes should be performed according to Clauses 4     to 9 of this standard.

  - Software development PROCESS

#### Software development planning

The objective of this ACTIVITY is to plan the software development  TASKS  to  reduce  RISKS  caused by software, communicate procedures and goals to members of the development team, and ensure that SYSTEM quality requirements for the MEDICAL DEVICE SOFTWARE are   met.

The software development planning ACTIVITY can document TASKS in a single plan or in multiple plans. Some MANUFACTURERS  might have established policies and procedures that apply to    the

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 49 –

+AMD1:2015 CSV  IEC 2015

development of all their MEDICAL DEVICE SOFTWARE. In this case the plan can simply reference the existing policies and procedures. Some MANUFACTURERS might prepare a  plan or  set of plans specific to the development of each MEDICAL DEVICE SOFTWARE PRODUCT that spell out in detail specific ACTIVITIES and reference general procedures. Another possibility is that a plan or set of plans is tailored for the development of each MEDICAL DEVICE SOFTWARE PRODUCT. The planning should be specified at the level of detail necessary to carry out the development PROCESS  and should be proportional to the RISK. For example, SYSTEMS  or  items with higher RISK would be subject to a development PROCESS with more rigor and TASKS should be spelled out in greater detail.

Planning is an iterative ACTIVITY that should be re-examined and updated as development progresses. The plan can evolve to incorporate more and better information as more is understood about the SYSTEM and the level of effort needed to develop the  SYSTEM.  For  example, a SYSTEM’s initial software safety classification can change as a  result of  exercising the RISK MANAGEMENT PROCESS and development of the software ARCHITECTURE. Or it might be decided that a SOUP be incorporated into the SYSTEM. It is important that the plan(s) be updated to reflect current knowledge of  the SYSTEM  and the level of  rigor needed for the SYSTEM  or items in the SYSTEM to enable proper control over the development   PROCESS.

#### Software requirements analysis

This ACTIVITY requires the MANUFACTURER to establish and verify the software requirements for the MEDICAL DEVICE SOFTWARE. Establishing verifiable requirements is essential for determining what is to be built, for determining that the MEDICAL DEVICE SOFTWARE exhibits acceptable behaviour, and for demonstrating that the completed  MEDICAL  DEVICE  SOFTWARE  is  ready for use. To demonstrate that the requirements have been implemented as desired, each  requirement should be stated in such a way that objective criteria can be established to  determine whether it has been implemented correctly. If the device RISK MANAGEMENT PROCESS imposes requirements on the software to control identified RISKS, these requirements are to be identified in the software requirements in such a way as to make it possible to trace the RISK CONTROL measures to the software requirements. All software requirements should  be  identified in such a way as to make it possible to demonstrate TRACEABILITY between the requirement and SOFTWARE SYSTEM testing. If regulatory approval in some countries requires conformance to specific regulations or international standards, this conformance requirement should be documented in the software requirements. Because the software requirements establish what is to be implemented in the software, an evaluation of the  requirements  is required before the requirements analysis ACTIVITY is  complete.

An area of frequent confusion is the distinction between customer needs, design  inputs,  software requirements, software functional specifications, and software design specifications. Design inputs are the interpretation of customer needs into formally  documented  MEDICAL  DEVICE requirements. Software requirements are the formally  documented  specifications  of what the software does to meet the customer needs and the design inputs. Software functional specifications are often included with the software requirements and define in detail what the software does to meet its requirements even  though  many different alternatives  might  also meet the requirements. Software design specifications define how the software will  be  designed and decomposed to implement its requirements and functional  specifications.

Traditionally, software requirements, functional specifications, and design specifications have been written as a set of one or more documents. It is now feasible to consider this information   as data items within a common database. Each item would have one or more attributes that would define its purpose and linkage to other items in the database. This approach allows presentation and printing of different views of the information best suited for each  set  of intended users (e.g., marketing, MANUFACTURERS, testers, auditors) and supports TRACEABILITY  to demonstrate adequate implementation and the extent to which test cases test the requirements. Tools to support this approach can be as simple as a hypertext document using HTML hyperlinks or as complex and capable as computer aided software engineering (CASE) tools and requirements analysis tools.

The SYSTEM requirements PROCESS is out of scope of this standard. However, the decision to implement MEDICAL  DEVICE  functionality with software is  normally made during SYSTEM    design.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 50 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Some or all of the SYSTEM requirements are allocated to be implemented in software.  The software requirements analysis ACTIVITY consists of analyzing the requirements allocated to software by the SYSTEM requirements PROCESS and deriving a comprehensive set of software requirements that reflect the allocated requirements.

To ensure the integrity of the SYSTEM, the MANUFACTURER should provide a mechanism for negotiating changes and clarifications to the SYSTEM requirements to correct impracticalities, inconsistencies or ambiguities in either the parent SYSTEM requirements or the software requirements.

The PROCESS of capture and analysis of SYSTEM and software requirements can be  iterative.  This standard does not intend to require the PROCESSES to  be  rigidly  segregated  into  two layers. In practice, SYSTEM ARCHITECTURE and software ARCHITECTURE are often outlined simultaneously and the SYSTEM and software requirements are subsequently documented in a layered form.

    - Software ARCHITECTURAL design

This ACTIVITY requires the MANUFACTURER to define the major structural components of the software and identify their key responsibilities, their externally visible properties, and the relationship among them. If the behaviour of a component can affect other components, that behavior should be described in the software ARCHITECTURE. This description is especially important for behaviour that can affect components of the MEDICAL DEVICE that are outside the software (see 5.3.5 and B.4.3). ARCHITECTURAL decisions are extremely important for implementing RISK CONTROL measures. Without understanding (and documenting) the behaviour of a  component that can affect other components, it  will be  nearly impossible to  show that the SYSTEM is safe. A software ARCHITECTURE is necessary to ensure the correct implementation of the software requirements. The software ARCHITECTURE is not  complete  unless all software requirements can be implemented by the identified SOFTWARE  ITEMS.  Because the design and implementation of the software is dependent on the ARCHITECTURE, the ARCHITECTURE is VERIFIED to complete this ACTIVITY. VERIFICATION of the ARCHITECTURE  is generally done by technical EVALUATION.

The software safety classification of SOFTWARE ITEMS during the  software  ARCHITECTURE  ACTIVITY creates a basis for the subsequent choice of software PROCESSES. The records of classification are placed under change control as part of the RISK MANAGEMENT   FILE.

Many subsequent events might invalidate the classification. These include, for  example:

- changes of SYSTEM specification, software specification or  ARCHITECTURE;

- discovery of errors in the RISK ANALYSIS, especially unforeseen HAZARDS;  and

- discovery of the infeasibility of a requirement, especially a RISK CONTROL   measure;

Therefore, during all ACTIVITIES following the design of the software ARCHITECTURE, the classification of the SOFTWARE SYSTEM and SOFTWARE ITEMS should be re-EVALUATED and might need to be revised. This would trigger rework to apply additional PROCESSES  to a  SOFTWARE   ITEM as a result of its upgrading to a higher class. The software configuration management PROCESS (Clause 8) is used to ensure that all necessary rework is identified and   completed.

#### Software detailed design

This ACTIVITY requires the MANUFACTURER to refine the SOFTWARE ITEMS  and interfaces defined  in the ARCHITECTURE to create SOFTWARE UNITS  and their interfaces. Although SOFTWARE UNITS  are often thought of as being a  single function or module, this view is not always appropriate. We have This standard has defined SOFTWARE UNIT to be a SOFTWARE ITEM that  is  not  subdivided into smaller items. SOFTWARE UNITS can be tested separately. The MANUFACTURER should define the level of detail of the SOFTWARE  UNIT. Detailed design specifies algorithms,  data representations, interfaces among different SOFTWARE UNITS, and interfaces between SOFTWARE UNITS and data structures. Detailed design must also be concerned with  the packaging of the SOFTWARE PRODUCT. It is necessary to document the design  of each  SOFTWARE UNIT  and its interface so that the SOFTWARE UNIT  can be implemented correctly.    The

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 51 –

+AMD1:2015 CSV  IEC 2015

detailed design fills in the details necessary to construct the software. It is necessary to define the design of the SOFTWARE UNITS and the interfaces in sufficient detail to permit its SAFETY and effectiveness to be objectively VERIFIED where this can be ensured using other requirements or design documentation. It should be complete enough that the programmer is not required to make ad hoc design decisions. Detailed design must also be concerned with the architecture of the MEDICAL DEVICE  SOFTWARE.

A SOFTWARE ITEM can be decomposed so that only a few of  the  new  SOFTWARE  ITEMS  implement the SAFETY-related requirement of the original SOFTWARE ITEM. The remaining SOFTWARE ITEMS do not implement SAFETY-related functions and  can  be  reclassified  into  a lower software safety class. However, the decision to do this is in itself part of the RISK MANAGEMENT PROCESS, and is documented in the RISK MANAGEMENT   FILE.

Because implementation depends on detailed design, it is necessary to verify the  detailed  design before the ACTIVITY is complete. VERIFICATION of detailed design is generally done by a technical EVALUATION. Subclause 5.4.4 requires the MANUFACTURER to verify the outputs of the detailed design ACTIVITIES. The design specifies how the requirements are to be implemented. VERIFICATION of the design provides assurance that it implements the software ARCHITECTURE  and is free from contradiction with the software  ARCHITECTURE.

If the design contains defects, the code will not implement the requirements  correctly.

When present in the design, the MANUFACTURER should verify design characteristics which the

MANUFACTURER believes are important for SAFETY. Examples of these characteristics  include:

- implementation of the intended events, inputs, outputs, interfaces, logic flow, allocation of CPU, allocation of memory resources, error and exception definition, error and exception isolation, and error recovery;

- definition of the default state, in which all faults that can result in a hazardous situation are addressed, with events and transitions;

- initialization of variables, memory management; and

- cold and warm resets, standby, and other state changes that can affect the RISK CONTROL

measures.

    - SOFTWARE UNIT implementation and  verification

This ACTIVITY requires the MANUFACTURER  to write and verify the code for the SOFTWARE UNITS.  The detailed design is to be translated into source code. Coding represents the point where decomposition of  the specifications ends and composition of the executable software begins.  To consistently achieve the desirable code characteristics, coding standards should be used to specify a preferred coding style. Examples of coding standards include requirements for understandability, language usage rules or restrictions, and complexity management. The code for each unit is VERIFIED to ensure that it functions as specified by the detailed design and that     it complies with the specified coding standards.

Subclause 5.5.5 requires the MANUFACTURER to verify the code. If the code does not implement  the design correctly, the MEDICAL DEVICE SOFTWARE will not perform as   intended.

#### Software integration and integration testing

This ACTIVITY  requires the MANUFACTURER  to plan and execute integration of  SOFTWARE  UNITS   into aggregate SOFTWARE ITEMS as well as integration of SOFTWARE ITEMS  into  higher  aggregated SOFTWARE ITEMS and to verify that the resulting SOFTWARE ITEMS behave  as  intended.

The approach to integration can range from non-incremental integration to any form of incremental integration. The properties of the SOFTWARE ITEM being assembled dictate  the chosen method of integration.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 52 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Software integration testing focuses on the transfer of data and control across a  SOFTWARE ITEM’s internal and external interfaces. External interfaces are those with other software, including operating system software, and MEDICAL DEVICE  hardware.

The rigor of integration testing and the level of detail of the documentation associated with integration testing should be commensurate with the RISK associated with the device, the device’s dependence on software for potentially hazardous functions, and the role of specific SOFTWARE ITEMS in higher RISK device functions. For example, although all SOFTWARE ITEMS should be tested, items that have an effect on SAFETY should be subject to more  direct,  thorough, and detailed tests.

As applicable, integration testing demonstrates program behaviour  at  the  boundaries  of  its input and output domains and confirms program responses to invalid, unexpected, and special inputs. The program’s actions are revealed when given combinations of inputs or unexpected sequences of inputs, or when defined timing requirements are violated. The test requirements    in the plan should include, as  appropriate, the types of  white box testing to be performed as part of integration testing.

White box testing, also known as glass box, structural, clear box and open box testing, is a testing technique where explicit knowledge of the internal workings of the SOFTWARE ITEM being tested are used to select the test data. White box testing uses specific knowledge of the SOFTWARE ITEM to examine outputs. The test is accurate only if the tester knows what the SOFTWARE ITEM  is supposed to do. The tester can then see if the SOFTWARE ITEM diverges from  its intended goal. White box testing cannot guarantee that the complete specification has been implemented since it is focused on testing the implementation of the SOFTWARE ITEM. Black box testing, also known as behavioural, functional, opaque-box, and closed-box testing, is focused  on testing the functional specification and it cannot guarantee that all parts of the implementation have been tested. Thus black  box testing is testing against the specification  and will discover faults of omission, indicating that part of the specification has not  been  fulfilled.  White  box  testing  is testing  against  the  implementation  and   will   discover  faults of commission, indicating that part of the implementation is faulty. In order to fully test a MEDICAL DEVICE SOFTWARE PRODUCT both black and white box testing might be   required.

The plans and test documentation identified in 5.6 and 5.7 can be individual documents tied to specific phases of development or evolutionary prototypes. They also might be combined so a single document or set of documents covers the requirements of multiple subsections. All or portions of the documents could be incorporated into higher level project documents such as a software or project quality assurance plan or a comprehensive test plan that addresses all aspects of testing for hardware and software. In these cases, a cross reference should be created that identifies how the various project documents relate to each of the software integration TASKS.

Software integration testing can be performed in a simulated environment, on actual target hardware, or on the full MEDICAL  DEVICE.

Subclause 5.6.2 requires the MANUFACTURER to verify the output of the software integration ACTIVITY. The output of the software integration ACTIVITY is the integrated SOFTWARE ITEMS. These integrated SOFTWARE ITEMS must function properly for the entire MEDICAL  DEVICE  SOFTWARE to function correctly and safely.

    - SOFTWARE SYSTEM testing

This ACTIVITY requires the MANUFACTURER  to verify the software’s functionality by verifying that  the requirements for the software have been successfully  implemented.

SOFTWARE SYSTEM testing demonstrates that the specified functionality exists. This testing VERIFIES the functionality and performance of the program as built with respect to the requirements for the software.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 53 –

+AMD1:2015 CSV  IEC 2015

SOFTWARE SYSTEM testing focuses on functional (black box) testing, although it might be  desirable to use white box (see previous section) methods to  more  efficiently  accomplish certain tests, initiate stress conditions or faults, or increase code coverage of the qualification tests. The organization of testing by types and test stage is flexible, but coverage of requirements, RISK CONTROL, usability, and test types (e.g., fault, installation, stress) should be demonstrated and documented.

SOFTWARE SYSTEM testing tests the integrated software and can be performed in a simulated environment, on actual target hardware, or on the full MEDICAL  DEVICE.

When a change is made to a SOFTWARE SYSTEM (even a small change), the degree of REGRESSION TESTING (not just the testing of the individual change) should be determined to ensure that no  unintended side effects have been introduced. This REGRESSION  TESTING  (and the rationale for not fully repeating SOFTWARE SYSTEM testing) should be planned and documented. (See B.6.3).

SOFTWARE SYSTEM test responsibilities can be dispersed, occurring at different locations and being conducted by different organizations. However, regardless of the distribution of TASKS, contractual relations, source of components, or development environment, the device MANUFACTURER retains ultimate responsibility for ensuring that the  software functions properly  for its intended use.

If  ANOMALIES  uncovered during testing can be  repeated, but a  decision has been made not to fix them, then these ANOMALIES need to be EVALUATED in relation to the HAZARD RISK analysis to verify that they do not affect the SAFETY of the device. The root cause and symptoms of the ANOMALIES should be understood, and the rationale for not fixing them should be   documented.

Subclause 5.7.4 requires the results of the SOFTWARE SYSTEM testing be EVALUATED to ensure  that the expected results were obtained.

#### Software release

This ACTIVITY requires the MANUFACTURER to document the VERSION of the MEDICAL DEVICE SOFTWARE being released, specify how it was created, and follow appropriate procedures for release of the software.

The MANUFACTURER should be able to show that the software that was developed using the development PROCESS is  the software that is  being released. The MANUFACTURER  should also  be able to retrieve the software and the tools used for its generation in case it is needed in the future and should store, package, and deliver the software in a manner that minimizes the software from being damaged or misused. Defined procedures should be established to ensure that these TASKS are performed appropriately and with consistent  results.

  - Software maintenance PROCESS

#### Establish software maintenance plan

The software maintenance PROCESS differs from the software  development  PROCESS  in  two ways:

- The MANUFACTURER is permitted to use a smaller PROCESS than the full  software  development PROCESS to implement rapid changes in response to urgent  problems.

- In responding to software PROBLEMS REPORTS relating to released product,  the MANUFACTURER not only addresses the problem but also satisfies local regulations (typically  by running a pro-active surveillance scheme for collecting problem data from the field and communicating with users and regulators about the  problem).

Subclause 6.1 requires these PROCESSES to be established in a maintenance   plan.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 54 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

This ACTIVITY requires the MANUFACTURER to create or identify procedures for implementing maintenance ACTIVITIES and TASKS. To implement corrective actions, control changes during maintenance, and manage release of revised software, the  MANUFACTURER  should document and resolve reported problems and requests from  users, as  well as  manage modifications to the MEDICAL DEVICE SOFTWARE. This PROCESS is activated when the MEDICAL DEVICE SOFTWARE undergoes modifications to code and associated documentation because of either a problem or the need for improvement or adaptation. The objective is to modify released MEDICAL DEVICE SOFTWARE while preserving its integrity. This PROCESS  includes  migration  of  the  MEDICAL DEVICE SOFTWARE to environments or platforms for which it was not originally released. The ACTIVITIES provided in this clause are specific to the maintenance PROCESS; however, the maintenance PROCESS might use other PROCESSES in this  standard.

The MANUFACTURER  needs  to plan how the ACTIVITIES  and TASKS  of the maintenance    PROCESS

will be performed.

#### Problem and modification analysis

This ACTIVITY requires the MANUFACTURER to analyze feedback for its effect; verify reported problems; and consider, select, and obtain approval for implementing a modification option. Problems and other requests for changes can affect the performance, SAFETY, or regulatory clearance of a MEDICAL DEVICE. An analysis is necessary to determine whether any effects exist because of a PROBLEM REPORT or whether any effects will result from a modification to correct a problem or implement a request. It is especially important to verify through trace or regression analysis that the RISK CONTROL measures built into the device are not adversely changed or modified by the software change that is being implemented as part  of the  software  maintenance ACTIVITY. It is also important to verify that the modified software does not cause a HAZARD HAZARDOUS SITUATION or mitigate a RISK in software that previously did not cause  a HAZARD HAZARDOUS SITUATION or mitigate RISKS. The software safety  classification  of a SOFTWARE ITEM might have changed if the software modification now can cause a HAZARD or mitigate a RISK.

It is important to distinguish between software maintenance (Clause 6) and software problem resolution (Clause 9).

The focus of the software maintenance PROCESS is an adequate response to feedback arising after release of the MEDICAL DEVICE SOFTWARE PRODUCT. As part of a MEDICAL DEVICE,  the software maintenance PROCESS needs to ensure  that:

- SAFETY-related PROBLEM REPORTS are addressed and reported to appropriate regulatory authorities and affected users;

- MEDICAL DEVICE SOFTWARE PRODUCTS are is re-validated and re-released after modification with formal controls that ensure the rectification of the problem and the avoidance of further problems;

- the MANUFACTURER considers what other MEDICAL DEVICE SOFTWARE PRODUCTS might be affected and takes appropriate action.

The focus of software problem resolution is the operation of a comprehensive control system that:

- analyses PROBLEM REPORTS and identifies all the implications of the   problem;

- decides on a number of changes and identifies all their  side-effects;

- implements the changes while maintaining the consistency of the software CONFIGURATION ITEMS  including the RISK MANAGEMENT FILE;

- VERIFIES the implementation of the  changes.

The software maintenance PROCESS uses the software problem resolution PROCESS. The  software maintenance PROCESS handles the high-level decisions about the PROBLEM REPORT (whether a problem exists, whether it has a significant effect on SAFETY, what changes are needed and when to implement them), and uses the software problem resolution PROCESS to analyse the PROBLEM REPORT to discover all the implications and to generate possible CHANGE

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 55 –

+AMD1:2015 CSV  IEC 2015

REQUESTS  which  identify  all  the  CONFIGURATION  ITEMS  that  need  to  be  changed  and  all the

VERIFICATION steps that are  necessary.

#### Modification implementation

This ACTIVITY requires that the MANUFACTURER use an established PROCESS to make the modification. If a maintenance PROCESS has not been defined, the appropriate development PROCESS TASKS can be used to make the modification. The MANUFACTURER should also ensure that the modification does not cause a negative effect on other parts of the MEDICAL DEVICE SOFTWARE. Unless the MEDICAL DEVICE SOFTWARE is treated as a new development, analysis of the effect of a modification on the entire MEDICAL DEVICE SOFTWARE is necessary. Regression analysis and testing are employed to provide assurance that a change has  not  created problems elsewhere in the MEDICAL DEVICE SOFTWARE. Regression analysis is the determination of the impact of a change based on review of the relevant documentation (e.g., software requirements specification, software design specification, source code, test plans, test cases, test scripts, etc.) in order to identify the necessary regression tests to  be  run.  Regression testing is the rerunning of test cases that a program has previously executed correctly and comparing the current result to the previous result in order to detect unintended effects of a software change. A rationale must be made that justifies the amount of  REGRESSION  TESTING  that will be performed to ensure that the portions of the MEDICAL DEVICE SOFTWARE not being modified still perform as they did before the modification was  made.

B.7     Software RISK MANAGEMENT  PROCESS

Software RISK MANAGEMENT is a part of overall MEDICAL DEVICE RISK MANAGEMENT and cannot be adequately addressed in isolation. This standard requires the use of a RISK MANAGEMENT  PROCESS that is compliant with ISO 14971. RISK MANAGEMENT as defined in ISO 14971 deals specifically with a framework for effective management of the RISKS associated with the use of MEDICAL DEVICES. One portion of ISO 14971 pertains to  control of  identified RISKS  associated with each HAZARD identified during the RISK ANALYSIS. The software RISK  MANAGEMENT  PROCESS  in this standard is intended to provide additional requirements for RISK CONTROL for software, including software that has been identified during the RISK  ANALYSIS  as potentially contributing  to a hazardous situation, or software that is used to control MEDICAL DEVICE RISKS. The software RISK MANAGEMENT PROCESS is included in this standard for two   reasons.

- the intended audience of this standard needs to understand minimum requirements for RISK CONTROL measures in their area of  responsibility—software;

- the general RISK MANAGEMENT standard, ISO 14971, provided as  a  normative reference in this standard, does not specifically address the RISK CONTROL of software  and  the  placement of RISK CONTROL in the software development life  cycle.

Software RISK MANAGEMENT is a part of overall MEDICAL DEVICE RISK MANAGEMENT. Plans, procedures, and documentation required for the software RISK MANAGEMENT ACTIVITIES can be a series of separate documents or a single document, or they can be integrated with the MEDICAL DEVICE RISK MANAGEMENT ACTIVITIES and documentation as long as all requirements in this standard are met.

#### B.7.1       Analysis of software contributing to hazardous situations

It is expected that the device HAZARD analysis will identify hazardous situations  and corresponding RISK CONTROL measures to reduce the probability and/or severity of those hazardous situations to an acceptable level. It is also expected  that  the  RISK  CONTROL  measures will be assigned to software functions that are expected to implement those RISK CONTROL measures.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 56 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

However, it is not expected that all device hazardous situations can be identified until the software ARCHITECTURE has been produced. At that time it is known how software functions will  be implemented in software components, and the practicality of the RISK CONTROL measures assigned to software functions can be EVALUATED. At that time the  device  HAZARD  analysis should be revised to include:

- revised hazardous situations;

- revised RISK CONTROL measures and software  requirements;

- new hazardous situations arising from  software, for example hazardous situations related    to human factors.

The software ARCHITECTURE should include credible strategies for segregating software components so that they do not interact in unsafe  ways.

### Software configuration management PROCESS

The software configuration management PROCESS is a PROCESS of applying administrative and technical procedures throughout the software life cycle to identify and define SOFTWARE ITEMS, including documentation, in a SYSTEM; control modifications and releases of the items; and document and report the status of the items and CHANGE REQUESTS. Software configuration management is necessary to recreate a SOFTWARE ITEM, to identify its constituent parts, and to provide a history of the changes that have been made to  it.

#### Configuration identification

This ACTIVITY requires the MANUFACTURER to uniquely identify software CONFIGURATION ITEMS and their VERSIONS. This identification is necessary to identify the software  CONFIGURATION  ITEMS  and the VERSIONS that are included in the MEDICAL DEVICE   SOFTWARE.

#### Change control

This ACTIVITY requires the MANUFACTURER to control changes of the software CONFIGURATION ITEMS and to document information identifying CHANGE REQUESTS and providing documentation about their disposition. This ACTIVITY is necessary to ensure that unauthorized or unintended changes are not made to the software CONFIGURATION ITEMS and to ensure  that  approved CHANGE REQUESTS are implemented fully and  verified.

CHANGE REQUESTS can be approved by a change control board or by a  manager or technical lead according to the software configuration management plan.  Approved  CHANGE  REQUESTS are made traceable to the actual modification and VERIFICATION of the  software.  The requirement is that each actual change be linked to a CHANGE REQUEST and that documentation exists to show that the CHANGE REQUEST was approved. The documentation might be change control board minutes, an approval signature, or a record in a  database.

#### Configuration status accounting

This ACTIVITY requires the MANUFACTURER to maintain records of the history of the software CONFIGURATION ITEMS. This ACTIVITY is necessary to determine when and why changes were made. Access to this information is necessary to ensure that software CONFIGURATION ITEMS contain only authorized modifications.

### B.9     Software problem resolution PROCESS

The software problem resolution PROCESS is a PROCESS for analyzing and resolving  the  problems (including non-conformances), whatever their nature or source, including those discovered during the execution of development, maintenance, or other PROCESSES. The objective is to provide a timely, responsible, and documented means to ensure that discovered problems  are  analyzed  and  resolved  and  that  trends  are  recognized.  This        PROCESS   is

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 57 –

+AMD1:2015 CSV  IEC 2015

sometimes called “defect tracking” in software engineering literature. It is called “problem resolution” in ISO/IEC 12207 [9] and IEC 60601-1-4 [2], Amendment 1. We have chosen to call    it “software problem resolution” in this  standard.

This ACTIVITY requires that the MANUFACTURER use the  software  problem  resolution  PROCESS when a problem or non-conformance is identified. This ACTIVITY is necessary to ensure that discovered problems are analyzed and EVALUATED for possible relevance to  SAFETY  (as specified in ISO 14971).

Software development plan(s) or procedures, as required in 5.1, are to address how problems    or  non-conformances will be handled. This includes specifying at each stage of the life cycle    the aspects of the software problem resolution PROCESS that will be formal and documented as well as when problems and nonconformities are to be entered into the software problem resolution PROCESS.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 58 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

# Annex C

## (informative)

Relationship to other standards

### General

This standard applies to the development and maintenance of MEDICAL DEVICE SOFTWARE. The software is considered a subsystem of the MEDICAL DEVICE or is itself a MEDICAL DEVICE. This standard is to be used together with other appropriate standards when developing a MEDICAL DEVICE.

MEDICAL DEVICE management standards such as ISO 13485 [8] (see C.2 and Annex D) and ISO 14971 (see Annex C.3) provide a management environment that lays a foundation for an organization to develop products. Safety standards such as  IEC 60601-1 [1] (see Annex C.4)  and IEC 61010-1 [5] (see Annex C.5) give specific direction for creating safe MEDICAL DEVICES. When software is a part of these MEDICAL DEVICES, IEC 62304 provides more detailed direction on what is required to develop and maintain safe MEDICAL DEVICE SOFTWARE. Many other standards such as ISO/IEC 12207 [9] (see Annex C.6), IEC 61508-3 [4] (see Annex C.7) and ISO/IEC 90003 [15] can be looked to as a source of methods, tools and techniques that can be used to implement the requirements in IEC 62304. Figure C.1 shows the relationship of these standards.

Where clauses or requirements from other standards are quoted, defined terms in the quoted items are terms that are defined in the other standard, not defined terms in this   standard.

Gives detailed direction

how to develop and maintain safe software system

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 59 –

Medical device management standards ISO 14971

ISO 13485

Medical device product standards

IEC 60601-1,

IEC 61010-1

Gives specific direction for creation of a safe medical device

Lays out a foundation to develop a medical device

affects

affects

requires

+AMD1:2015 CSV  IEC 2015

Medical device process standard IEC 62304

Implementation of medical device software

affects

inspires

Other sources of information IEC/ISO 12207,

IEC 61508-3,

IEC/ISO 90003, ...

Gives additional guidelines, techniques, etc that may be used

Medical device product standards

IEC 60601-1

IEC 61010-1

IEC 82304-1

Medical device management standards ISO 14971

ISO 13485

Gives specific direction for creation of a safe medical device

requires

IEC  725/06

Lays out a foundation to

develop a medical device

affects

affects

Medical device process standards

IEC 62304

IEC 62366-1

affects

Implementation of medical device software

inspires

Other sources of information IEC/ISO 12207

IEC 61508-3

IEC/ISO 90003, ...

Gives additional guidelines, techniques, etc that may be used

Gives detailed direction

how to develop and maintain safe software system

IEC

Figure C.1 – Relationship of key MEDICAL DEVICE standards to IEC  62304

### Relationship to ISO 13485

– 60 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

This standard requires that the MANUFACTURER  employs a quality management system.  When   a MANUFACTURER uses ISO 13485 [8], the requirements of ISO 62304 directly relate to some of the requirements of ISO 13485 as shown in Table  C.1.

#### Table C.1 – Relationship to ISO 13485:2003

| IEC 62304 clause | Related clause of ISO 13485:2003 |
| --- | --- |
| 5.1 Software development planning | 7.3.1 Design and development planning |
| 5.2 Software requirements analysis | 7.3.2 Design and development inputs |
| 5.3 Software ARCHITECTURAL design |  |
| 5.4 Software detailed design |  |
| 5.5 SOFTWARE UNIT implementation and verification |  |
| 5.6 Software integration and integration testing |  |
| 5.7 SOFTWARE SYSTEM testing | - Design and development outputs - Design and development review |
| 5.8 Software release | - Design and development verification - Design and development validation |
| 6.1 Establish software maintenance plan | 7.3.7 Control of design and development changes |
| 6.2 Problem and modification analysis |  |
| 6.3 Modification implementation | - Design and development verification - Design and development validation |
| 7.1 Analysis of software contributing to hazardous situations |  |
| 7.2 RISK CONTROL measures |  |
| 7.3 VERIFICATION of RISK CONTROL measures |  |
| 7.4 RISK MANAGEMENT of software changes |  |
| 8.1 Configuration identification | 7.5.3 Identification and TRACEABILITY |
| 8.2 Change control | 7.5.3 Identification and TRACEABILITY |
| 8.3 Configuration status accounting |  |
| 9 Software problem resolution PROCESS |  |

  - Relationship to ISO 14971

Table C.2 shows the areas where IEC 62304 amplifies requirements for the RISK MANAGEMENT PROCESS required by ISO 14971.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 61 –

+AMD1:2015 CSV  IEC 2015

#### Table C.2 – Relationship to ISO 14971:2000 2007

| ISO 14971:2000 2007 clause |  | Related clause of IEC 62304 |  |
| --- | --- | --- | --- |
| 4.1 | RISK ANALYSIS procedure process |  |  |
| 4.2 | Intended use/intended purpose and identification of characteristics related to the SAFETY of the MEDICAL DEVICE |  |  |
| 4.3 | Identification of known or foreseeable HAZARDS | 7.1 | Analysis of software contributing to HAZARDOUS SITUATIONS |
| 4.4 | Estimation of the RISK(S) for each HAZARD HAZARDOUS SITUATION | 4.3 | Software safety classification |
| 5 | RISK evaluation |  |  |
| 6.1 | RISK reduction |  |  |
| 6.2 | RISK CONTROL option analysis | 7.2.1 | Define RISK CONTROL measures |
| 6.3 | Implementation of RISK CONTROL measures | 7.2.2 7.3.1 | RISK CONTROL measures implemented in software Verify RISK CONTROL measures |
| 6.4 | RESIDUAL RISK evaluation |  |  |
| 6.5 | RISK/benefit analysis |  |  |
| 6.6 | Other generated HAZARDS RISKS arising from RISK CONTROL MEASURES | 7.3.2 | Document any new sequences of events |
| 6.7 | Completeness of RISK evaluation CONTROL |  |  |
| 7 | Evaluation of overall RESIDUAL RISK evaluation acceptability |  |  |
| 8 | RISK MANAGEMENT report | 7.3.3 | Document TRACEABILITY |
| 9 | Production and post-production information | 7.4 | RISK MANAGEMENT of software changes |

  - Relationship to PEMS requirements of IEC  60601-1:2005

+  IEC 606011:2005/AMD1:2012

    - General

Requirements for software are a subset of the requirements for a programmable electrical medical system (PEMS). This standard identifies requirements for software which are  in addition to, but not incompatible with, the requirements of IEC 60601-1:2005 + IEC 60601-  1:2005 /AMD1:2012 [1] for PEMS. Because PEMS include elements that are not software, not  all of the requirements of IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 for PEMS are addressed  in this  standard.  With  the  publication  of IEC 60601-1:2005 + IEC   60601-   1:2005 /AMD1:2012, IEC 62304 is now a normative reference of IEC 60601-1 and compliance with Clause 14 of IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 (and  thus  compliance  with  the  standard)  requires  compliance  with  parts  of IEC 62304  (not  with  the  whole  of IEC 62304 because IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 does not require compliance with post-production and maintenance requirements of IEC 62304). Finally, it is important to remember that IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 is only used if the software is part of a PEMS and not if the software is itself a MEDICAL DEVICE.

#### Software relationship to PEMS development

By using the V-model illustrated in Figure C.2 to describe what occurs during a PEMS development, it can be seen that the requirements of this software standard apply at the PEMS component level, from the specification of the software requirements to the integration of the SOFTWARE ITEMS into a SOFTWARE SYSTEM. This SOFTWARE SYSTEM is a part of a programmable electrical subsystem (PESS), which is a part of a  PEMS.

– 62 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| Software unit VERIFICATION (unit VERIFICATION) | Unit VERIFICATIO |
| --- | --- |
|  | results |

PEMS

requirements capture

PEMS validation

Subsystem (e.g. PESS) integration &

VERIFICATION

Software unit implementation

PEMS

architectural design

PEMS integration &

VERIFICATION

Subsystem (e.g.

PESS)

architectural design

User needs

Validated PEMS

PEMS validation plan

PEMS

validation results

PEMS requirement

specifications

Verified PEMS

PEMS test specification

PEMS VERIFICATION Plan

PEMS

VERIFICATION

results

PEMS architecture specification,

Subsystem (e.g. PESS) requirements  specifications

Verified Subsystem

Subsystem test specification

Subsystem

VERIFICATION

results

Verified software subsystem (component)

Software test specifications

Software  ARCHITECTURE specification

Verified code

N

Software integration & SOFTWARE SYSTEM VERIFICATION

(component integration & verification)

Software architectural design (component design)

Portion of PEMS V-model included in IEC 62304

Software requirements specifications (component  requirements)

Software integration and

VERIFICATION

results

Software detailed design

(unit design)

Key:

Boxes represent typical development lifecycle activities

Solid Arrows indicate typical deliverables transfered into/out of activities Dotted arrows indicate deliverables just to the Risk Management File

Outputs from problem resolution process Inputs to problem resolution process

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC  726/06

#### Figure C.2 – Software as part of the  V-model

    - Development PROCESS

Compliance with the software development PROCESS of this standard (Clause 5) requires that a software development plan be specified and then followed; it does not require  that  any particular life cycle model is  used, but it  does require that the plan include certain ACTIVITIES   and  have  certain  attributes.  These  requirements  relate  to   the   PEMS  requirements  in   IEC 60601-1 for development life cycle, requirement specification, ARCHITECTURE, design and implementation, and VERIFICATION. The requirements in this standard provide greater detail  about software development than those in IEC  60601-1.

    - Maintenance PROCESS

Compliance with the software maintenance PROCESS of this standard (Clause 6) requires that procedures be established and followed when changes to software are made. These require- ments correspond to the requirement in IEC 60601-1 for modification of a PEMS. The requirements in this standard for software maintenance  provide  greater  detail  about  what  must be done for software  maintenance  than  the  requirements  for  PEMS  modification  in IEC 60601-1.

    - Other PROCESSES

The other PROCESSES in this standard specify additional requirements for software beyond the similar requirements for PEMS in IEC 60601-1. In most cases, there is a general requirement for PEMS in IEC 60601-1, which the PROCESSES in this standard expand  upon.

The software RISK MANAGEMENT PROCESS in this standard corresponds to the additional RISK MANAGEMENT requirements identified for PEMS in IEC  60601-1.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 63 –

+AMD1:2015 CSV  IEC 2015

The software problem resolution PROCESS in this standard corresponds to the  problem  resolution requirement for PEMS in IEC 60601-1.

The software configuration management PROCESS in this standard specifies additional requirements that are not present for PEMS in IEC 60601-1 except for  documentation.

#### Coverage of PEMS requirements in IEC 60601-1:2005

+ IEC 606011:2005 /AMD1:2012

Table C.3 shows the PEMS requirements of IEC 60601-1 and the corresponding requirements  in this standard.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 64 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

#### Table C.3 – Relationship to IEC 60601-1

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.1 General The requirements of this clause in 14.2 to 14.12 (inclusive) shall apply to PEMS unless: - the PESS none of the PROGRAMMABLE ELECTRONIC SUBSYSTEMS (PESS) provides functionality necessary for no BASIC SAFETY or ESSENTIAL PERFORMANCE; or - the application of ISO 14971 the application of RISK MANAGEMENT as described in 4.2 demonstrates that the failure of the PESS does not lead to an unacceptable RISK. The requirements in 14.13 are applicable to any PEMS intended to be incorporated into an IT- NETWORK whether or not the requirements in 14.2 to 14.12 apply. When the requirements in 14.2 to 14.13 apply, the requirements in subclause 4.3, Clause 5, Clause 7, Clause 8 and Clause 9 of IEC 62304:2006 shall also apply to the development or modification of software for each PESS. | 4.3 Software safety classification The PEMS requirements of IEC 60601-1 would only apply to software safety classes B and C. This standard includes some requirements for software safety class A. The software development PROCESS required for compliance with IEC 60601-1 does not include the post production monitoring and maintenance required by Clause 6 of IEC 62304:2006. |
| 14.2 Documentation In addition to the records and documents required by ISO 14971, the documents produced from application of Clause 14 shall be maintained and shall form part of the RISK MANAGEMENT FILE. The documents required by Clause 14 shall be reviewed, approved, issued and changed in accordance with a formal document control procedure. | 4.2 RISK MANAGEMENT |
|  | 5.1 Software development planning In addition to the specific requirements in the software development planning ACTIVITY, documents that are part of the RISK MANAGEMENT FILE are required to be maintained by ISO 14971. In addition, for documents that are required by the quality system, ISO 13485 [8] requires control of the documents. |
| 14.3 RISK MANAGEMENT PLAN The RISK MANAGEMENT plan required by 3.5 of ISO 14971 4.2.2 shall also include a reference to the PEMS VALIDATION plan (see 14.11). | Not specifically required. There is no specific software validation plan. The PEMS validation plan is at the SYSTEM level and thus is outside the scope of this software standard. This standard does require TRACEABILITY from HAZARD to specific software cause to RISK CONTROL measure to VERIFICATION of the RISK CONTROL measure (see 7.3) |
| 14.4 PEMS DEVELOPMENT LIFE-CYCLE A PEMS DEVELOPMENT LIFE-CYCLE shall be documented. | 5.1 Software development planning 5.1.1 Software development plan The items addressed by the software development plan constitute a SOFTWARE DEVELOPMENT LIFE CYCLE. |
| The PEMS DEVELOPMENT LIFE-CYCLE shall contain a set of defined milestones. |  |
| At each milestone, the ACTIVITIES to be completed and the VERIFICATION methods to be applied to those activities shall be defined. | 5.1.6 Software VERIFICATION planning VERIFICATION TASKS, milestones and acceptance criteria must be planned. |
| Each activity shall be defined including its inputs and outputs. | 5.1.1 Software development plan ACTIVITIES are defined in this standard. Documentation to be produced is defined in each ACTIVITY. |
| Each milestone shall identify the RISK MANAGEMENT ACTIVITIES that must be completed before that milestone. |  |
| The PEMS DEVELOPMENT LIFE-CYCLE shall be tailored for a specific development by making plans which detail ACTIVITIES, milestones and schedules. | 5.1.1 Software development plan This standard allows the development life cycle to be documented in the development plan. This means the development plan contains a tailored development life cycle. |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 65 –

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| The PEMS DEVELOPMENT LIFE-CYCLE shall include documentation requirements. | 5.1.1 Software development plan 5.1.8 Documentation planning |
| 14.5 Problem resolution Where appropriate, a documented system for problem resolution within and between all phases and ACTIVITIES of the PEMS DEVELOPMENT LIFE- CYCLE shall be developed and maintained. | 9 Software problem resolution PROCESS |
| Depending on the type of product, the problem resolution SYSTEM may: - be documented as a part of the PEMS DEVELOPMENT LIFE-CYCLE; - allow the reporting of potential or existing problems affecting BASIC SAFETY or ESSENTIAL PERFORMANCE; - include an assessment of each problem for associated RISKS; - identify the criteria that must be met for the issue to be closed; - identify the action to be taken to resolve each problem. | 5.1.1 Software development plan 9.1 Prepare PROBLEM REPORTS |
| 14.6 RISK MANAGEMENT PROCESS | 7 Software RISK MANAGEMENT PROCESS |
| 14.6.1 Identification of known and foreseeable HAZARDS When compiling the list of known or foreseeable HAZARDS, the MANUFACTURER shall consider those HAZARDS associated with software and hardware aspects of the PEMS including those associated with NETWORK/DATA COUPLING the incorporation of the PEMS into an IT-NETWORK, components of third-party origin and legacy subsystems. | 7.1 Analysis of software contributing to HAZARDOUS SITUATIONS This standard does not mention network/data coupling specifically |
| 14.6.2 RISK CONTROL Suitably validated tools and PROCEDURES shall be selected and identified to implement each RISK CONTROL measure. These tools and PROCEDURES shall be appropriate to assure that each RISK CONTROL measure satisfactorily reduces the identified RISK(S). | 5.1.4 Software development standards, methods and tools planning This standard requires the identification of specific tools and methods to be used for development in general, not for each RISK CONTROL measure. |
| 14.7 Requirements specification For the PEMS and each of its subsystems (e.g. for a PESS) there shall be a documented requirement specification. | 5.2 Software requirements analysis This standard deals only with the software subsystems of a PEMS. |
| The requirement specification for a system or subsystem shall include and distinguish any ESSENTIAL PERFORMANCE and any RISK CONTROL measures implemented by that system or subsystem. | - Define and document software requirements from SYSTEM requirements. - Software requirements content - Include RISK CONTROL measures in software requirements This standard does not require that the requirements related to essential performance and RISK CONTROL measures be distinguished from other requirements, but it does require that all requirements be uniquely identified. |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 66 –	IEC 62304:2006

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.8 ARCHITECTURE For the PEMS and each of its subsystems, an ARCHITECTURE shall be specified that shall satisfy the requirements specification. | 5.3 Software ARCHITECTURAL design |
| Where appropriate, to reduce the RISK to an acceptable level, the architecture specification shall make use of: - COMPONENTS WITH HIGH-INTEGRITY CHARACTERISTICS; - fail-safe functions; - redundancy; - diversity; - partitioning of functionality; - defensive design, e.g. limits on potentially hazardous effects by restricting the available output power or by introducing means to limit the travel of actuators. | 5.3.5 Identify segregation necessary for RISK CONTROL Partitioning is the only technique identified, and it is only identified because there is a requirement to state how the integrity of the partitioning is assured. |
| The ARCHITECTURE specification shall take into consideration: - allocation of RISK CONTROL measures to subsystems and components of the PEMS; - failure modes of components and their effects; - common cause failures; - systemic failures; - test interval duration and diagnostic coverage; - maintainability; - protection from reasonably foreseeable misuse; - the IT-NETWORK/DATA COUPLING specification, if applicable. | This is not included in this standard. |
| 14.9 Design and implementation Where appropriate, the design shall be decomposed into subsystems, each having both a design and test specification. | 5.4 Software detailed design 5.4.2 Develop detailed design for each SOFTWARE UNIT This standard does not require a test specification for detailed design. |
| Descriptive data regarding the design environment shall be included in the RISK MANAGEMENT FILE documentation. | 5.4.2 Develop detailed design for each SOFTWARE UNIT |
| 14.10 VERIFICATION VERIFICATION is required for all functions that implement BASIC SAFETY, ESSENTIAL PERFORMANCE or RISK CONTROL measures. | 5.1.6 Software VERIFICATION planning VERIFICATION is required for each ACTIVITY |

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 67 –

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| A VERIFICATION plan shall be produced to show how these functions shall be verified. The plan shall include: - at which milestone(s) VERIFICATION is to be performed on each function; - the selection and documentation of VERIFICATION strategies, ACTIVITIES, techniques, and the appropriate level of independence of the personnel performing the VERIFICATION; - the selection and utilization of VERIFICATION tools; - coverage criteria for VERIFICATION. | 5.1.6 Software VERIFICATION planning Independence of personnel is not included in this standard. It is considered covered in ISO 13485. |
| The VERIFICATION shall be performed according to the VERIFICATION plan. The results of the VERIFICATION activities shall be documented. | VERIFICATION requirements are in most of the ACTIVITIES. |
| 14.11 PEMS VALIDATION A PEMS VALIDATION plan shall include the validation of BASIC SAFETY and ESSENTIAL PERFORMANCE, and shall require checks for unintended functioning of the PEMS. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| Methods used for PEMS VALIDATION shall be documented | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| The PEMS VALIDATION shall be performed according to the PEMS VALIDATION plan. The results of the PEMS VALIDATION activities shall be documented. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| The person having the overall responsibility for the PEMS VALIDATION shall be independent of the design team. The MANUFACTURER shall document the rationale for the level of independence. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| No member of a design team shall be responsible for the PEMS VALIDATION of their own design. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| All professional relationships of the members of the PEMS VALIDATION team with members of the design team shall be documented in the RISK MANAGEMENT FILE. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| A reference to the methods and results of the PEMS VALIDATION shall be included in the RISK MANAGEMENT FILE. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| 14.12 Modification If any or all of a design results from a modification of an earlier design then either all of this clause applies as if it were a new design or the continued validity of any previous design documentation shall be assessed under a documented modification/change PROCEDURE. | 6 Software maintenance PROCESS This standard takes the approach that software maintenance should be planned and that implementation of modifications should use the software development PROCESS or an established software maintenance PROCESS. |
| When software is modified, the requirements in subclause 4.3, Clause 5, Clause 7, Clause 8 and Clause 9 of IEC 62304:2006 shall also apply to the modification. |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 68 –	IEC 62304:2006

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.13 Connection of PEMS by NETWORK/DATA COUPLING to other equipment PEMS intended to be incorporated into an IT-NETWORK If the PEMS is intended to be connected by NETWORK/DATA COUPLING to other equipment that is outside the control of the PEMS MANUFACTURER, the technical description shall: a) specify the characteristics of the NETWORK/DATA COUPLING necessary for the PEMS to achieve its INTENDED USE; If the PEMS is intended to be incorporated into an IT-NETWORK that is not validated by the PEMS MANUFACTURER, the MANUFACTURER shall make available instructions for implementing such connection including the following - the purpose of the PEMS’S connection to an IT- NETWORK; - the required characteristics of the IT-NETWORK incorporating the PEMS; - the required configuration of the IT-NETWORK incorporating the PEMS; - the technical specifications of the network connection of the PEMS including security specifications; - the intended information flow between the PEMS, the IT-NETWORK and other devices on the IT-NETWORK, and the intended routing through the IT-NETWORK; and NOTE 1 This can include aspects of effectiveness and data and system security as related to BASIC SAFETY and ESSENTIAL PERFORMANCE (see also Clause H.6 and IEC 80001-1:2010). b)f) list the HAZARDOUS SITUATIONS resulting from a failure of the IT-NETWORK/DATA COUPLING to provide the specified characteristics required to meet the purpose of the PEMS connection to the IT-NETWORK. c) Instruct the RESPONSIBLE ORGANIZATION that: In the ACCOMPANYING DOCUMENTS, the MANUFACTURER shall instruct the RESPONSIBLE ORGANIZATION that: - connection of the PEMS to a NETWORK/DATA COUPLING an IT-NETWORK that includes other equipment could result in previously unidentified RISKS to PATIENTS, OPERATORS or third parties; - the RESPONSIBLE ORGANIZATION should identify, analyze, evaluate and control these RISKS; | Requirements for network/data coupling incorporation into an IT-network are not included in this standard. |

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 69 –

+AMD1:2015 CSV  IEC  2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| NOTE 3 IEC 80001-1:2010 provides guidance for the RESPONSIBLE ORGANIZATION to address these risks. - subsequent changes to the IT-NETWORK/DATA COUPLING could introduce new RISKS and require additional analysis; and - changes to the IT-NETWORK/DATA COUPLING include: - changes in the IT-NETWORK/DATA COUPLING configuration; - connection of additional items to the IT- NETWORK/DATA COUPLING; - disconnecting items from the IT- NETWORK/DATA COUPLING; - update of equipment connected to the IT- NETWORK/DATA COUPLING; - upgrade of equipment connected to the IT-NETWORK/DATA COUPLING. |  |

#### C.4.7	Relationship to requirements in IEC  60601-1-4

IEC 60601-1-4 will continue to be used until the transition period for IEC 60601-1:2005 is complete.

Table C.4 shows the requirements of IEC 60601-1-4 [2] and the related requirements in this standard. This does not indicate that the related requirements in this standard fully cover the requirements in IEC 60601-1-4. Many parts of the 60601-1-4 requirements are covered by compliance with ISO 14971. Some  requirements  in  IEC 60601-1-4  are  not  addressed  by IEC 62304.

#### Table C.4 – Relationship to IEC  60601-1-4

| PEMS requirements from IEC 60601-1-4:1996 plus Amendment 1:1999 | Related requirements of IEC 62304 |
| --- | --- |
| 6.8 Accompanying documents |  |
| 6.8.201 | 4.2 and 4.3 c) |
| 52.201 Documentation |  |
| 52.201.1 | 4.1 |
| 52.201.2 | 4.1 and 4.2 |
| 52.201.3 | 4.2 |
| 52.202 RISK MANAGEMENT PLAN |  |
| 52.202.1 | 4.2 |
| 52.202.2 | 5.1.1, 5.1.5 |
| 52.202.3 | 4.1, 5.1.2 |
| 52.203 Development life-cycle |  |
| 52.203.1 | 5.1.1 |
| 52.203.2 | 5.1.1 |
| 52.203.3 |  |
| 52.203.4 | 5.1.7 |
| 52.203.5 | 7 |
| 52.204 Risk management process |  |
| 52.204.1 | 4.2 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 70 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1-4:1996 plus Amendment 1:1999 | Related requirements of IEC 62304 |
| --- | --- |
| 52.204.2 | 4.2, 7 |
| 52.204.3 |  |
| 52.204.3.1 |  |
| 52.204.3.1.1 | 4.2, 7.1 |
| 52.204.3.1.2 | 4.2, 7.1.2 |
| 52.204.3.1.3 | 4.2 |
| 52.204.3.1.4 | 4.2, 7.1.2 e) |
| 52.204.3.1.5 | 4.2, 7.1.2 |
| 52.204.3.1.6 | 4.2, 7.1 |
| 52.204.3.1.7 | 4.2 |
| 52.204.3.1.8 | 4.2 |
| 52.204.3.1.9 | 4.2 |
| 52.204.3.1.10 | 4.2 |
| 52.204.3.2 |  |
| 52.204.3.2.1 | 4.2 |
| 52.204.3.2.2 | 4.2, 4.3 |
| 52.204.3.2.3 |  |
| 52.204.3.2.4 |  |
| 52.204.3.2.5 | 4.2 |
| 52.204.4 |  |
| 52.204.4.1 | 4.2 |
| 52.204.4.2 | 4.2 |
| 52.204.4.3 | 4.2 |
| 52.204.4.4 | 4.2 |
| 52.204.4.5 |  |
| 52.204.4.6 | 4.2 |
| 52.205 Qualification of personnel | 4.1 |
| 52.206 Requirement specification |  |
| 52.206.1 | 5.2 |
| 52.206.2 | 7.2.2 |
| 52.206.3 |  |
| 52.207 Architecture |  |
| 52.207.1 | 5.3.1 |
| 52.207.2 | 5.3 |
| 52.207.3 |  |
| 52.207.4 |  |
| 52.207.5 |  |
| 52.208 Design and implementation |  |
| 52.208.1 | 5 |
| 52.208.2 |  |
| 52.209 Verification |  |
| 52.209.1 | 5.7.1 |
| 52.209.2 | 5.1.5, 5.1.6 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 71 –

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1-4:1996 plus Amendment 1:1999 | Related requirements of IEC 62304 |
| --- | --- |
| 52.209.3 | 5.2.6, 5.3.6, 5.4.4, 5.5.5, 5.6, 5.7 |
| 52.209.4 |  |
| 52.210 Validation |  |
| 52.210.1 | 4.1 |
| 52.210.2 | 4.1 |
| 52.210.3 | 4.1 |
| 52.210.4 |  |
| 52.210.5 |  |
| 52.210.6 |  |
| 52.210.7 |  |
| 52.211 Modification |  |
| 52.211.1 | 6 |
| 52.211.2 | 4.1, 6 |
| 52.212 Assessment |  |
| 52.212.1 | 4.1 |

IEC 60601-1-4 has been withdrawn.

### Relationship to IEC 61010-1

The scope of IEC 61010-1 [4 5] covers electrical test and measuring  equipment,  electrical control equipment and electrical laboratory equipment. Only part of the laboratory equipment is used in a medical environment or as in vitro diagnostic equipment  (IVD).

Due to legal regulations or normative references, IVD equipment is allocated  to  MEDICAL  DEVICES without, however, falling within the scope of IEC 60601-1 [1]. This is attributable not  only to the fact that, strictly speaking, IVD instruments are not  MEDICAL  DEVICES  which  come into direct contact with patients, but also to the fact that such products are manufactured for many different applications in various laboratories. Use as an IVD instrument or  as  an accessory for an IVD instrument is then rare.

If laboratory equipment is used as IVD equipment, the measured results obtained must be EVALUATED in accordance with medical criteria. The application of ISO  14971 is  required for  RISK MANAGEMENT. If such products also contain software that can lead to a HAZARD HAZARDOUS SITUATION, for example failure caused by the software which results in an unwanted change of medical data (measuring results), IEC 62304 must be taken into  account.

IEC 61010-1:2010 has a general requirement for risk assessment in Clause 17, which is more streamlined than the full risk management requirements of ISO 14971. Applying IEC 61010-1 Clause 17 alone does not meet the required criteria for risk management of IEC 62304, which     is based on full ISO 14971 risk management requirements. With this in mind, it is expected by this standard that when an IVD medical device has software-related risks, its risk management process is performed following ISO 14971 instead of only Clause 17 of IEC 61010-1. Compliance with Clause 17 of IEC 61010-1 will be achieved, as detailed in the Note to Clause    17 of IEC 61010-1:

NOTE  One RISK  assessment procedure is outlined in Annex J. Other RISK  assessment procedures are contained in ISO 14971, SEMI S10-1296, IEC 61508, ISO 14121-1, and ANSI B11.TR3. Other established procedures which implement similar steps can also be  used.

The flowchart in Figure C.3 provides a useful aid to explain the principle way of the RISK MANAGEMENT PROCESS and shows the application of IEC 62304 with IEC 61010-1, Clause   17:

– 72 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Intended purpose and use defined

Possible sources of HAZARD

identified

HAZARD related to the handling of medical data

Identify known and reasonably foreseeable

HAZARDS

Is the HAZARD covered by relevant safety standards?

Yes

Verify according to the relevant safety standard

No

Use

Yes

Does the device	No

ISO 14971

for RISK MANAGEMENT

provide medical relevant data

Does

the software have	No

any impact on the medical

data?

Select an applicable method for RISK CONTROL

based on safety standard

Additional requirements necessary to ensure that wrong data are detected prior to use of data for medical purposes.

Yes

Use of procedures required to verify the data?

Yes

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

No

Use IEC 62304

IEC  727/06

#### Figure C.3 – Application of IEC 62304 with IEC  61010-1

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 73 –

+AMD1:2015 CSV  IEC 2015

### Relationship to ISO/IEC 12207

This standard has been derived from the approach and concepts of ISO/IEC 12207 [9], which defines requirements for software life cycle PROCESSES in general, i.e. not restricted to MEDICAL DEVICES.

This standard differs from ISO/IEC 12207 mainly with respect to the following.  It:

- excludes SYSTEM aspects, such as SYSTEM requirements, SYSTEM ARCHITECTURE and validation;

- omits some PROCESSES seen as duplicating ACTIVITIES documented elsewhere for MEDICAL DEVICES;

- adds the (SAFETY) RISK MANAGEMENT PROCESS  and the software release  PROCESS;

- incorporates the documentation and the VERIFICATION supporting PROCESSES into the development and maintenance PROCESSES;

- merges the PROCESS  implementation and planning ACTIVITIES of each PROCESS  into a  single

ACTIVITY in the development and maintenance  PROCESSES;

- classifies the requirements with respect to SAFETY needs;  and

- does not explicitly classify PROCESSES as primary or supporting, nor group PROCESSES as ISO/IEC 12207 does.

Most  of  these changes  were  driven by the  desire to  tailor  the  standard to  the need of      the

MEDICAL DEVICE  sector by:

- focusing on SAFETY aspects and the MEDICAL DEVICE RISK MANAGEMENT standard ISO    14971;

- selecting the appropriate PROCESSES useful in a regulated  environment;

- taking into account that software development is embedded in a  quality system  (which covers some of the PROCESSES and requirements of ISO/IEC 12207);  and

- lowering the level of abstraction to make it easier to  use.

This standard is not contradictory to ISO/IEC 12207. ISO/IEC 12207 can be useful as an aide    in setting up a well structured SOFTWARE DEVELOPMENT LIFE CYCLE MODEL that includes the requirements of this standard.

Table C.5, which  was prepared  by  ISO/IEC  JTC1/SC7,  shows  the  relationship  between  IEC 62304 and ISO/IEC 12207.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 74 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

#### Table C.5 – Relationship to ISO/IEC 12207:2008

| ISO/IEC 62304 processes |  | ISO/IEC 12207 processes |  |
| --- | --- | --- | --- |
| Activity | Task | Activity | Task |
| 5 Software development PROCESS |  | 5.3 Development process - Documentation process - Configuration management process 6.4 Verification process 6.5 Validation process 6.8 Problem resolution process 7.1 Management process |  |
| 5.1 Software development planning |  | 5.3.1 Process implementation 5.3.3 System architectural design - Software coding and testing - Software integration - Software qualification testing - System integration 6.1.1 Process implementation 6.2.1 Process implementation 6.2.2 Configuration identification 6.4.1 Process implementation 6.5.1 Process implementation 6.8.1 Process implementation 7.1.2 Planning 7.1.3 Execution and control - Establishment of the infrastructure - Maintenance of the infrastructure |  |
|  | 5.1.1 Software development plan | 5.3.1 Process implementation | 5.3.1.1 |
|  |  | 7.1.2 Planning | 5.3.1.3 |
|  |  |  | 5.3.1.4 |
|  |  |  | 7.1.2.1 |
|  | 5.1.2 Keep software development plan updated | 7.1.3 Execution and control | 7.1.3.3 |
|  | 5.1.3 Software development plan | 5.3.3 System architectural | 5.3.3.1 |
|  | reference to SYSTEM design and development | design 5.3.10 System integration | 5.3.10.1 6.5.1.4 |
|  |  | 6.5.1 Process implementation |  |
|  | 5.1.4 Software development | 5.3.1 Process implementation | 5.3.1.3 |
|  | standards, methods and tools planning |  | 5.3.1.4 |
|  | 5.1.5 Software integration and integration testing planning | 5.3.8 Software integration. | 5.3.8.1 |
|  | 5.1.6 Software VERIFICATION | 6.4.1 Process implementation | 6.4.1.4 |
|  | planning | 5.3.7 Software coding and | 6.4.1.5 |
|  |  | testing | 5.3.7.5 |
|  |  | 5.3.8 Software integration | 5.3.8.5 |
|  |  | 5.3.9 Software qualification testing | 5.3.9.3 |
|  | 5.1.7 Software RISK MANAGEMENT planning | Amd.1:2002 – F 3.1.5 Risk management process |  |
|  | 5.1.8 Documentation planning | 6.1.1 Process implementation | 6.1.1.1 |
|  | 5.1.9 Software configuration | 6.2.1 Process implementation | 6.2.1.1 |
|  | management planning | 6.8.1 Process implementation | 6.8.1.1 |
|  | 5.1.10 Supporting items to be | 7.2.2 Establishment of the | 7.2.2.1 |
|  | controlled | infrastructure |  |
|  |  | 7.2.3 Maintenance of the infrastructure | 7.2.3.1 |
|  | 5.1.11 Software CONFIGURATION ITEM control before VERIFICATION | 6.2.2 Configuration identification | 6.2.2.1 |
| 5.2 Software |  | 5.3.3 System architectural |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 75 –

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 processes |  | ISO/IEC 12207 processes |  |
| --- | --- | --- | --- |
| Activity | Task | Activity | Task |
| requirements analysis |  | design 5.3.4 Software requirements analysis 6.4.2 Verification |  |
|  | 5.2.1 Define and document software requirements from SYSTEM requirements | 5.3.3 System architectural design | 5.3.3.1 |
|  | 5.2.2 Software requirements content | 5.3.4 Software requirements analysis | 5.3.4.1 |
|  | 5.2.3 Include RISK CONTROL measures in software requirements |  |  |
|  | 5.2.4 Re-EVALUATE MEDICAL DEVICE RISK ANALYSIS |  | None |
|  | 5.2.5 Update SYSTEM requirements | 5.3.4 Software requirements analysis | a) b) |
|  | 5.2.6 Verify software requirements | 5.3.4 Software requirements analysis 6.4.2 Verification | 5.3.4.2 6.4.2.3 |
| 5.3 Software ARCHITECTURAL design |  | 5.3.5 Software architectural design |  |
|  | 5.3.1 Transform software requirements into an ARCHITECTURE | 5.3.5 Software architectural design | 5.3.5.1 |
|  | 5.3.2 Develop an ARCHITECTURE for the interfaces of SOFTWARE ITEMS |  | 5.3.5.2 |
|  | 5.3.3 Specify functional and performance requirements of SOUP item |  | none |
|  | 5.3.4 Specify SYSTEM hardware and software required by SOUP item |  | none |
|  | 5.3.5 Identify segregation necessary for RISK CONTROL |  | none |
|  | 5.3.6 Verify software ARCHITECTURE | 5.3.5 Software architectural design | 5.3.5.6 |
| 5.4 Software detailed design |  | 5.3.6 Software detailed design 6.4.2 Verification |  |
|  | 5.4.1 Refine SOFTWARE ARCHITECTURE into SOFTWARE UNITS | 5.3.6 Software detailed design | 5.3.6.1 |
|  | 5.4.2 Develop detailed design for each SOFTWARE UNIT |  |  |
|  | 5.4.3 Develop detailed design for interfaces |  | 5.3.6.2 |
|  | 5.4.4 Verify detailed design | 6.4.2 Verification | 5.3.6.7 |
| 5.5 SOFTWARE UNIT implementation and verification |  | 5.3.6 Software detailed design 5.3.7 Software coding and testing 6.4.2 Verification |  |
|  | 5.5.1 Implement each SOFTWARE UNIT | 5.3.7 Software coding and testing | 5.3.7.1 |
|  | 5.5.2 Establish SOFTWARE UNIT VERIFICATION PROCESS | 5.3.6 Software detailed design 5.3.7 Software coding and testing | 5.3.6.5 5.3.7.5 |
|  | 5.5.3 SOFTWARE UNIT acceptance criteria | 5.3.7 Software coding and testing | 5.3.7.5 |
|  | 5.5.4 Additional SOFTWARE UNIT acceptance criteria | 5.3.7 Software coding and testing 6.4.2 Verification | 5.3.7.5 6.4.2.5 |
|  | 5.5.5 SOFTWARE UNIT VERIFICATION | 5.3.7 Software coding and testing | 5.3.7.2 |
| 5.6 Software integration and integration testing |  | 5.3.8 Software integration 5.3.9 Software qualification |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 76 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 processes |  | ISO/IEC 12207 processes |  |
| --- | --- | --- | --- |
| Activity | Task | Activity | Task |
|  |  | testing 5.3.10 System integration 6.4.1 Process implementation 6.4.2 Verification |  |
|  | 5.6.1 Integrate SOFTWARE UNITS | 5.3.8 Software integration | 5.3.8.2 |
|  | 5.6.2 Verify software integration | 5.3.8 Software integration 5.3.10 System integration | 5.3.8.2 5.3.10.1 |
|  | 5.6.3 Test integrated software | 5.3.9 Software qualification testing. | 5.3.9.1 |
|  | 5.6.4 Integration testing content |  | 5.3.9.3 |
|  | 5.6.5 Verify integration tests procedures | 6.4.2 Verification | 6.4.2.2 |
|  | 5.6.6 Conduct regression tests | 5.3.8 Software integration | 5.3.8.2 |
|  | 5.6.7 Integration test record contents | 5.3.8 Software integration | 5.3.8.2 |
|  | 5.6.8 Use software problem resolution PROCESS | 6.4.1 Process implementation | 6.4.1.6 |
| 5.7 SOFTWARE SYSTEM testing |  | - Software integration - Software qualification testing 6.4.1 Process implementation 6.4.2 Verification 6.8.1 Process implementation |  |
|  | 5.7.1 Establish tests for each software requirement | - Software integration - Software qualification testing | 5.3.8.4 5.3.9.1 |
|  | 5.7.2 Use software problem resolution PROCESS | 6.4.1 Process implementation | 6.4.1.6 |
|  | 5.7.3 Retest after changes | 6.8.1 Process implementation | 6.8.1.1 |
|  | 5.7.4 Verify SOFTWARE SYSTEM testing | 6.4.2 Verification 5.3.9 Software qualification testing | 6.4.2.2 5.3.9.3 |
|  | 5.7.5 Document data for each test SOFTWARE SYSTEM test record content | 5.3.9 Software qualification testing | 5.3.9.1 |
| 5.8 Software release |  | 5.3.9 Software qualification testing 5.4.2 Operational testing - Configuration evaluation - Release management and delivery |  |
|  | 5.8.1 Ensure software VERIFICATION is complete | 5.4.2 Operational testing 6.2.6 Release management and delivery | 5.4.2.1 5.4.2.2 6.2.6.1 |
|  | 5.8.2 Document known residual ANOMALIES | 6.2.5 Configuration evaluation 5.3.9 Software qualification testing | 6.2.5.1 5.3.9.3 |
|  | 5.8.3 Evaluate known residual ANOMALIES |  |  |
|  | 5.8.4 Document released VERSIONS | 6.2.6 Release management and delivery | 6.2.6.1 |
|  | 5.8.5 Document how released software was created |  |  |
|  | 5.8.6 Ensure activities and tasks are complete |  |  |
|  | 5.8.7 Archive software |  |  |
|  | 5.8.8 Assure repeatability of software release |  |  |
| 6 Software maintenance PROCESS |  | 5.5 Maintenance process 6.2 Configuration management process |  |
| 6.1 Establish software maintenance plan |  | 5.5.1 Process implementation | 5.5.1.1 |
| 6.2 Problem and modification analysis |  | 5.5.1 Process Implementation 5.5.2 Problem and modification |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 77 –

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 processes |  | ISO/IEC 12207 processes |  |
| --- | --- | --- | --- |
| Activity | Task | Activity | Task |
|  |  | analysis 5.5.3 Modification implementation 5.5.5 Migration |  |
|  | 6.2.1 Record and evaluate feedback |  |  |
|  | 6.2.1.1 Monitor feedback | 5.5.1 Process Implementation | 5.5.1.1 5.5.1.2 |
|  | 6.2.1.2 Document and EVALUATE feedback |  |  |
|  | 6.2.1.3 Evaluate PROBLEM REPORT’S affects on SAFETY | 5.5.2 Problem and modification analysis | 5.5.2.1 5.5.2.2 5.5.2.3 5.5.2.4 |
|  | 6.2.2 Use software problem resolution PROCESS | 5.5.1 Process Implementation | 5.5.1.2 |
|  | 6.2.3 Analyse CHANGE REQUESTS | 5.5.2 Problem and modification analysis | 5.5.2.1 |
|  | 6.2.4 CHANGE REQUEST approval | 5.5.2 Problem and modification analysis | 5.5.2.5 |
|  | 6.2.5 Communicate to users and regulators | 5.5.3 Modification implementation 5.5.5 Migration | 5.5.3.1 5.5.5.3 |
| 6.3 Modification implementation |  | 5.5.3 Modification implementation 6.2.6 Release management and delivery |  |
|  | 6.3.1 Use established PROCESS to implement modification | 5.5.3 Modification implementation | 5.5.3.2 |
|  | 6.3.2 Re-release modified SOFTWARE SYSTEM | 6.2.6 Release management and delivery | 6.2.6.1 |
| 7 Software RISK MANAGEMENT PROCESS |  | Amd.1:2002 – F 3.15 Risk management process Process in 62304 addresses risk / hazard issues that are not addressed in Amd 1. There is some commonality (risk measures, etc) but the focus of the analysis is quite different. |  |
| 8 Software configuration management PROCESS |  | 5.5 Maintenance process 6.2 Configuration management process |  |
| 8.1 Configuration identification |  | 6.2.2 Configuration identification |  |
|  | 8.1.1 Establish means to identify CONFIGURATION ITEMS | 6.2.2 Configuration identification | 6.2.2.1 |
|  | 8.1.2 Identify SOUP |  | none |
|  | 8.1.3 Identify SYSTEM configuration documentation | 6.2.2 Configuration identification | 6.2.2.1 |
| 8.2 Change control |  | 5.5.3 Modification implementation 6.2.3 Configuration control |  |
|  | 8.2.1 Approve CHANGE REQUESTS | 6.2.3 Configuration control | 6.2.3.1 |
|  | 8.2.2 Implement changes | 5.5.3 Modification implementation 6.2.3 Configuration control | 5.5.3.2 6.2.3.1 |
|  | 8.2.3 Verify changes | 6.2.3 Configuration control | 6.2.3.1 |
|  | 8.2.4 Provide means for TRACEABILITY of change |  |  |
| 8.3 Configuration status accounting |  | 6.2.4 Configuration status accounting | 6.2.4.1 |
| 9 Software problem resolution PROCESS |  | 5.5 Maintenance process 6.2 Configuration management 6.8 Problem resolution process |  |
| 9.1 Prepare PROBLEM REPORTS |  | 6.8.1 Process implementation 6.8.2 Problem resolution | 6.8.1.1 b) 6.8.2.1 |
| 9.2 Investigate the problem |  | 6.8.2 Problem resolution 6.8.1 Process implementation | 6.8.2.1 6.8.1.1 b) |
| 9.3 Advise relevant parties |  | 6.8.1 Process implementation | 6.8.1.1 a) |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 78 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 processes |  | ISO/IEC 12207 processes |  |
| --- | --- | --- | --- |
| Activity | Task | Activity | Task |
| 9.4 Use change control process |  | 6.2.3 Configuration control. 5.5.3 Modification implementation |  |
| 9.5 Maintain records |  | 6.8.1 Process implementation | 6.8.1.1 a) |
| 9.6 Analyse problems for trends |  | 6.8.1 Process implementation 6.8.2 Problem resolution | 6.8.1.1 b) 6.8.2.1 |
| 9.7 Verify software problem resolution |  | 6.8.1 Process implementation | 6.8.1.1 d) |
| 9.8 Test documentation contents |  |  | All testing tasks in 12207 require documentation |

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
| 5 Software development PROCESS |  |  |  |
| 5.1 Software development planning |  |  |  |
|  | 5.1.1 Software development plan | 7.1.1 Software Implementation | 7.1.1.3.1 Software implementation strategy 7.1.1.3.1.1 7.1.1.3.1.3 7.1.1.3.1.4 6.3.1.3.2 Project planning 6.3.1.3.2.1 |
|  | 5.1.2 Keep software development plan updated | 6.3.2 Project Assesment and Control | 6.3.2.3.2 Project control 6.3.2.3.2.1 |
|  | 5.1.3 Software development plan reference to SYSTEM design and development | 6.4.3 System Architectural Design 6.4.5 System Integration 7.2.5 Software Validation Process | 6.4.3.3.1 Establishing architecture 6.4.3.3.1.1 6.4.5.3.1 Integration 6.4.5.3.1.1 7.2.5.3.1 Process implementation 7.2.5.3.1.4 |
|  | 5.1.4 Software development standards, methods and tools planning | 7.1.1 Software Implementation | 7.1.1.3.1 Software implementation strategy 7.1.1.3.1.3 |
|  | 5.1.5 Software integration and integration testing planning | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.1 |
|  | 5.1.6 Software VERIFICATION planning | 7.2.4 Software Verification - Software Construction - Software Integration - Software Qualification Testing | 7.2.4.3.1 Process implementation 7.2.4.3.1.4 7.2.4.3.1.5 7.1.5.3.1 Software constrution 7.1.5.3.1.5 7.1.6.3.1 Software integration 7.1.6.3.1.5 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.1.7 Software RISK MANAGEMENT planning | 6.3.4 Risk Management Process |  |
|  | 5.1.8 Documentation planning | 7.2.1 Software Documentation Management | 7.2.1.3.1 Process implementation 7.2.1.3.1.1 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 79 –

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
|  | 5.1.9 Software configuration management planning | 7.2.2 Software Configuration Management 7.2.8 Software Problem Resolution | 7.2.2.3.1 Process implementation 7.2.2.3.1.1 7.2.8.3.1 Process implementation 7.2.8.3.1.1 |
|  | 5.1.10 Supporting items to be controlled | 6.2.2 Infrastructure Management 6.2.2 Infrastructure Management | - Establishment of the infrastructure 6.2.2.3.2.1 - Maintenance of the infrastructure 6.2.2.3.3.1 |
|  | 5.1.11 Software CONFIGURATION ITEM control before VERIFICATION | 7.2.2 Software Configuration Management | 7.2.2.3.2 Configuration identification 7.2.2.3.2.1 |
| 5.2 Software requirements analysis |  |  |  |
|  | 5.2.1 Define and document software requirements from SYSTEM requirements | 6.4.3 System Architectural Design | 6.4.3.3.1 Establishing architecture 6.4.3.3.1.1 |
|  | 5.2.2 Software requirements content | 7.1.2 Software Requirements Analysis | 7.1.2.3.1 Software requirements analysis 7.1.2.3.1.1 |
|  | 5.2.3 Include RISK CONTROL measures in software requirements |  |  |
|  | 5.2.4 Re-EVALUATE MEDICAL DEVICE RISK ANALYSIS | None | None |
|  | 5.2.5 Update SYSTEM requirements | 7.1.2 Software Requirements Analysis | 7.1.2.3.1 Software requirements analysis 7.1.2.3.1.1 a) & b) |
|  | 5.2.6 Verify software requirements | 7.2.4 Software Verification | 7.2.4.3.2 Verification 7.2.4.3.2.1 |
| 5.3 Software ARCHITECTURAL design | 5.3.1 Transform software requirements into an ARCHITECTURE | 7.1.3 Software Architectural Design | 7.1.3.3.1 Software architectural design 7.1.3.3.1.1 |
|  | 5.3.2 Develop an ARCHITECTURE for the interfaces of SOFTWARE ITEMS |  | 7.1.3.3.1 Software architectural design 7.1.3.3.1.2 |
|  | 5.3.3 Specify functional and performance requirements of SOUP item | None | none |
|  | 5.3.4 Specify SYSTEM hardware and software required by SOUP item | None | none |
|  | 5.3.5 Identify segregation necessary for RISK CONTROL | None | none |
|  | 5.3.6 Verify software ARCHITECTURE | 7.1.3 Software Architectural Design | 7.1.3.3.1 Software architectural design 7.1.3.3.1.6 |
| 5.4 Software detailed design | 5.4.1 Refine SOFTWARE ARCHITECTURE into SOFTWARE UNITS | 7.1.4 Software Detailed Design | 7.1.4.3.1 Software detailed design 7.1.4.3.1.1 |
|  | 5.4.2 Develop detailed design for each SOFTWARE UNIT |  |  |
|  | 5.4.3 Develop detailed design for interfaces |  | 7.1.4.3.1 Software detailed design 7.1.4.3.1.2 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 80 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
|  | 5.4.4 Verify detailed design | 7.1.4 Software Detailed Design | 7.1.4.3.1 Software detailed design 7.1.4.3.1.7 |
| 5.5 SOFTWARE UNIT implementation and verification | 5.5.1 Implement each SOFTWARE UNIT | 7.1.5 Software Construction | 7.1.5.3.1 Software construction 7.1.5.3.1.1 |
|  | 5.5.2 Establish SOFTWARE UNIT VERIFICATION PROCESS | - Software Detailed Design - Software Construction | 7.1.4.3.1 Software detailed design 7.1.4.3.1.5 7.1.5.3.1 Software Construction 7.1.5.3.1.5 |
|  | 5.5.3 SOFTWARE UNIT acceptance criteria | 7.1.5 Software Construction | 7.1.5.3.1 Software construction 7.1.5.3.1.5 |
|  | 5.5.4 Additional SOFTWARE UNIT acceptance criteria | 7.1.5 Software Construction 7.2.4 Software Verification | 7.1.5.3.1 Software construction 7.1.5.3.1.2 |
|  | 5.5.5 SOFTWARE UNIT VERIFICATION | 7.1.5 Software Construction | 7.1.5.3.1 Software construction 7.1.5.3.1.2 |
| 5.6 Software integration and integration testing | 5.6.1 Integrate SOFTWARE UNITS | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 |
|  | 5.6.2 Verify software integration | 7.1.6 Software Integration 6.4.5 System Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 6.4.5.3.1 Integration 6.4.5.3.1.2 |
|  | 5.6.3 Test integrated software | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.1 |
|  | 5.6.4 Integration testing content | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.6.5 Verify integration tests procedures | None | None |
|  | 5.6.6 Conduct regression tests | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 |
|  | 5.6.7 Integration test record contents | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 |
|  | 5.6.8 Use software problem resolution PROCESS | 7.2.4 Software Verification | 7.2.4.3.1 Process implementation 7.2.4.3.1.6 |
| 5.7 SOFTWARE SYSTEM testing | 5.7.1 Establish tests for each software requirement | - Software Integration - Software Qualification Testing | 7.1.6.3.1 Software integration 7.1.6.3.1.4 7.1.7.3.1 Software qualification testing 7.1.7.3.1.1 |
|  | 5.7.2 Use software problem resolution PROCESS | 7.2.4 Software Verification | 7.2.4.3.1 Process implementation 7.2.4.3.1.6 |
|  | 5.7.3 Retest after changes | 7.2.8 Software Problem Resolution | 7.2.8.3.1 Process implementation 7.2.8.3.1.1 |
|  | 5.7.4 Verify SOFTWARE SYSTEM testing | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 81 –

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
|  | 5.7.5 SOFTWARE SYSTEM test record contents | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.1 |
| 5.8 Software release | 5.8.1 Ensure software VERIFICATION is complete | 6.4.9 Software Operation 7.2.2 Software Configuration Management | 6.4.9.3.2 Operation activation and check-out 6.4.9.3.2.1 6.4.9.3.2.2 7.2.2.3.6 Release management and delivery 7.2.2.3.6.1 |
|  | 5.8.2 Document known residual ANOMALIES | 7.2.2 Software Configuration Management 7.1.7 Software Qualification Testing | 7.2.2.3.5 Configuration evaluation 7.2.2.3.5.1 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.8.3 EVALUATE known residual ANOMALIES |  |  |
|  | 5.8.4 Document released VERSIONS | 7.2.2 Software Configuration Management Process | 7.2.2.3.6 Release management and delivery 7.2.2.3.6.1 |
|  | 5.8.5 Document how released software was created |  |  |
|  | 5.8.6 Ensure ACTIVITIES and TASKS are complete |  |  |
|  | 5.8.7 Archive software |  |  |
|  | 5.8.8 Assure repeatability of software release |  |  |
| 6 Software maintenance PROCESS |  | 6.4.10 Software Maintenance Process |  |
| 6.1 Establish software maintenance plan |  | 6.4.10 Software Maintenance | None |
| 6.2 Problem and modification analysis | 6.2.1 Document and EVALUATE feedback | None | None |
|  | 6.2.1.1 Monitor feedback | 6.4.10 Software Maintenance | None |
|  | 6.2.1.2 Document and EVALUATE feedback |  |  |
|  | 6.2.1.3 EVALUATE PROBLEM REPORT’S effects on SAFETY | 6.4.10 Software Maintenance | None |
|  | 6.2.2 Use software problem resolution PROCESS | 6.4.10 Software Maintenance | None |
|  | 6.2.3 Analyse CHANGE REQUESTS | 6.4.10 Software Maintenance | None |
|  | 6.2.4 CHANGE REQUEST approval | 6.4.10 Software Maintenance | None |
|  | 6.2.5 Communicate to users and regulators | 6.4.10 Software Maintenance | None |
| 6.3 Modification implementation |  | None | None |
|  | 6.3.1 Use established PROCESS to implement modification | 6.4.10 Software Maintenance | None |
|  | 6.3.2 Re-release modified SOFTWARE SYSTEM | 7.2.2 Software Configuration Management | None |
| 7 Software RISK MANAGEMENT PROCESS |  | 6.3.4 Risk Management Process This is based on ISO/IEC 16085. While there is some commonality it does not address the specific requirements for medical device software development with regard to risk management |  |
| 8 Software configuration management PROCESS |  |  |  |
| 8.1 Configuration identification | 8.1.1 Establish means to identify CONFIGURATION ITEMS | 7.2.2 Software Configuration Management | None |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 82 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| ISO/IEC 62304 PROCESSES |  |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- | --- |
| ACTIVITY | TASK |  | PROCESSES | ACTIVITY/TASK |
|  | 8.1.2 Identify SOUP |  | None | None |
|  | 8.1.3 Identify SYSTEM configuration documentation |  | 7.2.2 Software Configuration Management | None |
| 8.2 Change control | 8.2.1 Approve CHANGE REQUESTS |  | 7.2.2 Software Configuration Management | None |
|  | 8.2.2 Implement changes |  | 6.4.10 Software Maintenance | None |
|  | 8.2.3 | Verify changes | 7.2.2 Software Configuration Management | None |
|  | 8.2.4 Provide means for TRACEABILITY of change |  |  |  |
| 8.3 Configuration status accounting |  |  | 7.2.2 Software Configuration Management | None |
| 9 Software problem resolution PROCESS |  |  |  |  |
| 9.1 Prepare PROBLEM REPORTS |  |  | 7.2.8 Software Problem Resolution | None |
| 9.2 Investigate the problem |  |  | 7.2.8 Software Problem Resolution | None |
| 9.3 Advise relevant parties |  |  | 7.2.8 Software Problem Resolution | None |
| 9.4 Use change control |  |  | 7.2.2 Software | None |
| process |  |  | Configuration |  |
|  |  |  | Management |  |
|  |  |  | 6.4.10 Software |  |
|  |  |  | Maintenance |  |
| 9.5 Maintain records |  |  | 7.2.8 Software Problem Resolution | None |
| 9.6 Analyse problems for trends |  |  | 7.2.8 Software Problem Resolution | None |
| 9.7 Verify software problem resolution |  |  | 7.2.8 Software Problem Resolution | None |
| 9.8 Test documentation contents |  |  | All testing TASKS in ISO 12207 require documentation | None |

### Relationship to IEC 61508

The question has been raised whether this standard, being concerned with the design of SAFETY-critical software, should follow the principles of  IEC 61508. The approach to safety in IEC 62304 is fundamentally different than the one in IEC 61508. IEC 62304 takes into account that the effectiveness of medical devices justifies residual risks related to their use.  The following explains the stance of this  standard.

IEC 61508 addresses 3 main issues:

- RISK MANAGEMENT life cycle and life cycle  PROCESSES;

- definition of Safety Integrity Levels;

- recommendation of techniques, tools and methods for software development and levels of independence of personnel responsible for performing different  TASKS.

Issue 1) is covered in this standard by a normative reference to ISO 14971 (the MEDICAL DEVICE sector standard for RISK MANAGEMENT). The effect of this reference is to adopt ISO 14971’s approach to RISK MANAGEMENT as an integral part of the software PROCESS for MEDICAL DEVICE SOFTWARE.

For issue 2), this standard takes a simpler approach than IEC 61508. The latter classifies software into 4 “Safety Integrity Levels” defined in terms of reliability objectives. The reliability objectives are identified after RISK ANALYSIS, which quantifies both the severity and  the probability of HARM caused by a failure of the  software.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 83 –

+AMD1:2015 CSV  IEC 2015

This standard simplifies issue 2) by disallowing defining the classification into 3 software safety classes based on the RISK caused by a failure consideration of probability of software failure  prior to classification. Classification into 3 software safety classes is based only on the severity of that HARM caused by a failure. After classification, different PROCESSES are required for different software safety classes: the intention is to further reduce the probability (and/or the severity) of failure of the software.

Issue 3)  is  not addressed by this  standard. Readers of  the standard are encouraged to use   IEC 61508 as a source for good software methods, techniques and tools, while recognising that other approaches, both present and future, can provide equally good results. This standard makes no recommendation concerning independence of people responsible for one software ACTIVITY  (for example VERIFICATION) from those responsible for another (for example design).     In  particular, this standard makes no requirement for an independent safety assessor, since   this is a matter for ISO 14971.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 84 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

# Annex D

## (informative)

Implementation

### Introduction

This annex gives an overview of how this standard can be implemented into MANUFACTURERS’ PROCESSES. It also considers that other standards like ISO 13485 [8] require adequate and comparable PROCESSES.

### Quality management system

For MANUFACTURERS of MEDICAL DEVICES, including MEDICAL DEVICE SOFTWARE  in the context of this standard, the establishment of a quality management system (QMS) is required in 4.1. This standard does not require that the QMS necessarily has to be  certified.

  - EVALUATE  quality management PROCESSES

It is recommended to EVALUATE how well the established and documented PROCESSES of  the QMS already cover the PROCESSES  of the software life cycle, by means of audits, inspections,  or analyses under the responsibility of the MANUFACTURER. Any identified gaps can be accommodated by extending the QM PROCESSES, or can be separately described. If the MANUFACTURER already has PROCESS descriptions available which regulate the development, VERIFICATION and validation of  software, then these should also be  EVALUATED  to  determine how well they agree with this standard.

  - Integrating requirements of this standard into the MANUFACTURER’S quality management PROCESSES

This standard can be implemented by adapting or extending the PROCESSES already installed in the QMS system, or integrating new PROCESSES. This standard does not specify how this is to  be done; the MANUFACTURER is free to do this in any suitable   way.

The MANUFACTURER is responsible for ensuring that the PROCESSES  described in  this standard are suitably put into action when the MEDICAL DEVICE SOFTWARE is developed by Original Equipment Manufacturers (OEM) or sub-contractors not having their own documented  QMS.

  - Checklist for small MANUFACTURERS without a certified  QMS

The MANUFACTURER should determine the highest software safety classification (A, B or C) of  the software.  Table D.1  lists  all  ACTIVITIES  described  in  this  standard.  The  reference  to ISO 13485 should help to define the place in the QMS. Based on the required software safety class, the MANUFACTURER should assess each required ACTIVITY against  the  existing PROCESSES. If the requirement is already covered, a reference to the relevant PROCESS descriptions should be given.

If there is discrepancy, an action is needed to improve the   PROCESS.

The list can also be used for an EVALUATION of the PROCESSES after the action has been performed.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 85 –

+AMD1:2015 CSV  IEC 2015

#### Table D.1 – Checklist for small companies without a certified  QMS

| ACTIVITY | Related clause of ISO 13485:2003 | Covered by existing procedure? | If yes: Reference | Actions to be taken |
| --- | --- | --- | --- | --- |
| 5.1 Software development planning | 7.3.1 Design and development planning | Yes/No |  |  |
| 5.2 Software requirements analysis | 7.3.2 Design and development inputs | Yes/No |  |  |
| 5.3 Software ARCHITECTURAL design |  | Yes/No |  |  |
| 5.4 Software detailed design |  | Yes/No |  |  |
| 5.5 SOFTWARE UNIT implementation and verification |  | Yes/No |  |  |
| 5.6 Software integration and integration testing |  | Yes/No |  |  |
| 5.7 SOFTWARE SYSTEM testing | - Design and development outputs - Design and development review | Yes/No |  |  |
| 5.8 Software release | - Design and development verification - Design and development validation | Yes/No |  |  |
| 6.1 Establish software maintenance plan | 7.3.7 Control of design and development changes | Yes/No |  |  |
| 6.2 Problem and modification analysis |  | Yes/No |  |  |
| 6.3 Modification implementation | - Design and development verification - Design and development validation | Yes/No |  |  |
| 7.1 Analysis of software contributing to hazardous situations |  | Yes/No |  |  |
| 7.2 RISK CONTROL measures |  | Yes/No |  |  |
| 7.3 VERIFICATION of RISK CONTROL measures |  | Yes/No |  |  |
| 7.4 RISK MANAGEMENT of software changes |  | Yes/No |  |  |
| 8.1 Configuration identification | 7.5.3 Identification and traceability | Yes/No |  |  |
| 8.2 Change control | 7.5.3 Identification and traceability | Yes/No |  |  |
| 8.3 Configuration status accounting |  | Yes/No |  |  |
| 9 Software problem resolution PROCESS |  | Yes/No |  |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 86 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

## Bibliography

[1]	IEC 60601-1:2005, Medical electrical equipment – Part 1: General requirements   for basic safety and essential performance

IEC 60601-1:2005/AMD1:2012

[2]	IEC 60601-1-4:1996,  Medical  electrical  equipment  –  Part  1:  General requirements for safety – 4.Collateral standard: Programmable electrical medical systems  (withdrawn)

IEC 60601-1-4:1996/AMD1:1999

[3]	IEC 60601-1-6,  Medical electrical equipment –  Part 1-6: General requirements   for basic safety and essential performance – Collateral standard:  Usability

[4]	IEC 61508-3,  Functional  safety  of  electrical/electronic/programmable  electronic safety-

related systems – Part 3: Software  requirements

[5]	IEC 61010-1:2001 2010,  Safety  requirements  for  electrical equipment for measurement, control, and laboratory use – Part 1: General  requirements

[6]	ISO 9000:2005, Quality management systems – Fundamentals and  vocabulary

[7]	ISO 9001:2000 2008, Quality management systems –  Requirements

[8]	ISO  13485:2003, Medical devices –  Quality management systems     –  Requirements for regulatory purposes

[9]	ISO/IEC 9126-1:2001, Software engineering — Product quality — Part 1: Quality  model

[9]	ISO/IEC  12207:1995 2008, Information technology Systems and software   engineering – Software life cycle processes

Amendment 1 (2002)

Amendment 2 (2004)

[10]	ISO/IEC	14764:1999,   Information	technology	–	Software	maintenance	Software Engineering – Software Life Cycle Processes –  Maintenance

[11]	ISO/IEC  15504-5:2012,  Information  technology  –  Process  assessment  –  Part    5:  An exemplar software life cycle process assessment model

[12]	ISO/IEC 25010:2011, Systems and software engineering –  System and   software Quality Requirements and Evaluation (SQuaRE) – System and software quality  models

[13]	ISO/IEC 33001:—2),  Information  technology  –   Process  assessment  –    Concepts  and terminology

[14]	ISO/IEC 33004:—2), Information technology  –  Process  assessment –   Requirements for process reference, process assessment and maturity  models

[15]	ISO/IEC  90003:2004  2014,  Software  engineering  –  Guidelines  for  the   application  of ISO 9001:2000 2008 to computer software

[16]	ISO/IEC Guide 51:1999 2014, Safety aspects – Guidelines for their inclusion in   standards

[17]	IEEE 610.12:1990, IEEE standard glossary of software engineering  terminology

[18]	IEEE 1044:1993 2009, IEEE standard classification for software  anomalies

- To be published.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 87 –

+AMD1:2015 CSV  IEC 2015

[19] U.S. Department Of Health and Human Services, Food  and  Drug  Administration,  Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices, May 11, 2005,

<://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/uc m089543.htm>

[20] U.S. Department Of Health and Human Services, Food and Drug Administration, General Principles of Software Validation; Final Guidance for Industry and FDA Staff, January 11, 2002,

<http://www.fda.gov/downloads/RegulatoryInformation/Guidances/ucm126955.pdf>

[21]	IEC 62366-1:2015,  Medical  devices  –  Part  1:  Application  of  usability    engineering  to medical devices

[22]	IEC 82304-1:—3),  Healthcare Software Systems – Part 1: General  requirements

- In preparation.

– 88 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

# Index of defined terms

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

ACTIVITY, 15, 17, 23,    25, 27, 31, 33, 43, 59, 65,

67, 69, 73, 79, 81, 83, 87, 89, 95, 113, 133,

145

Change control, 101 Change request, 61 Completion of, 49

Configuration identification, 101 Configuration management, 35

Configuration status accounting, 101 Definition, 19

Deliverable, 19

Design and maintenance, 11 Hazard identification, 11

Maintenance, 51

Mapping, 15

Modification implementation, 97 Planning, 83, 85

Problem and modification analysis, 95 Problem resolution, 31, 53, 103

Required, 15, 147 Requirements, 17

Requirements analysis, 39 Risk analysis, 55

Risk management, 33, 47, 59, 79, 81, 99 Software architectural design, 87

Software detailed design, 89 Software development, 11

Software integration, 93

Software integration and integration testing, 91

Software maintenance, 95 Software release, 95

Software requirements analysis, 85 Software system testing, 93

SOFTWARE UNIT implementation and verification, 89

Testing, 45, 47

Verification, 33

ANOMALY, 45, 47, 49, 55, 65, 93

Definition, 19

ARCHITECTURE, 39, 41, 73, 75, 79, 81, 83, 85,

87, 89, 99, 113, 133

Definition, 19

CHANGE REQUEST, 53, 61, 63, 65, 97,  101

Definition, 19

CONFIGURATION ITEM, 27, 35, 49, 59, 61, 97,  101

Definition, 19

SOUP, 31, 59

DELIVERABLE, 25, 31, 33

Definition, 19

EVALUATION, 41,  45,  49,  51, 53, 55, 57, 87,  89,

93, 95, 99, 147, 149

Re-, 39

HARM, 21, 23, 73, 81, 145

Definition, 21

HAZARD, 11,  23,  29,  57,  67,  69, 79, 83, 93, 97,

99, 129

Definition, 21

Unforeseen, 87

MANUFACTURER, 15, 21, 23, 25, 27, 29, 31, 33,

35, 37, 39, 41, 43, 45, 47, 49, 51, 53, 55, 57,

59, 61, 63, 65, 75, 77, 79, 81, 83, 85, 87, 89,

91, 93, 95, 97, 101, 103, 107,  147

Definition, 21

MEDICAL  DEVICE, 11, 17, 21, 27, 35, 39, 41,    55,

69, 75, 77, 79, 85, 87, 91, 93, 95, 97, 99,   105,

129, 133, 145, 147

Definition, 21

MEDICAL  DEVICE  SOFTWARE, 11,  13,  17,  27, 35,

37, 39, 51, 67, 73, 75, 77, 79, 83, 85, 91, 93,

95, 97, 101, 105, 145, 147

Change, 59

Definition, 21

PROBLEM REPORT, 51, 53, 61, 63, 65, 95,  97

Classification, 61

Definition, 21

PROCESS, 13, 15, 17, 23, 25, 27, 31, 67, 69,   73,

75, 79, 81, 85, 87, 89, 97, 101, 103, 113,   133,

145, 147

Acceptance, 61

Change control, 61, 63 Classification, 133

Configuration management, 51, 89, 113 Decision-making, 77

Definition, 23

Development, 27, 81, 95, 113

Existing, 31

Improvement, 149

Life cycle, 11, 133, 143

Maintenance, 51, 53, 113

Mapping, 15

Modification, 97

Omission of, 81

Output, 75

Physiological, 21

Problem resolution, 35, 45, 47, 51, 53, 63,  97,

101, 103, 113

Quality management, 147 Required, 15, 147

Requirements, 17, 29

Risk analysis, 73

Risk management, 11, 23, 29, 33, 51, 63,  79,

81, 85, 89, 99, 109, 113, 129, 133

Software, 79, 145

Software development, 11, 27, 31, 53, 73

Software maintenance, 11, 95, 97 Software release, 133

System requirements, 87 Verification, 27

REGRESSION TESTING, 45, 65, 93

Definition, 23

RISK, 23, 67, 75, 79, 81, 83, 85, 91, 97,  99

Definition, 23

Non-serious injury, 29

Reasonably foreseeable, 79 Risk control, 23

Serious injury, 29

SOUP, 33

IEC 62304:2006	– 89 –

+AMD1:2015 CSV  IEC 2015

Unacceptable, 11, 25, 49

RISK ANALYSIS, 39, 55, 67, 73, 79, 87, 99,  145

Definition, 23 RISK CONTROL

Activity, 11

Definition, 23

Hardware measure, 29

Measure, 29, 31, 37, 43, 45, 55, 57, 59,  79,

81, 85, 87, 89, 93, 97, 99

Requirements, 39, 41, 57, 99

Segregation, 41

RISK  MANAGEMENT,  11,  23,  29,  33,  47,  51, 53,

59, 63, 67, 75, 77, 79, 81, 85, 87, 89, 99,   109,

113, 129, 133, 145

Definition, 23

Medical device, 75 Report, 57

RISK  MANAGEMENT  FILE, 17,  29,  55,  57,  63, 87,

89, 97

Definition, 23

SAFETY, 11, 51, 63, 69, 77, 81, 89, 91, 93, 95,

97, 103, 133, 143

Definition, 25

SECURITY, 63

Definition, 25

Requirements, 37 SERIOUS INJURY, 29, 83

Definition, 25

Non-, 29, 83

SOFTWARE DEVELOPMENT LIFE CYCLE MODEL, 31, 73, 133

Definition, 25

SOFTWARE  ITEM, 25, 27, 29, 31, 33, 39, 41,    43,

53, 55, 57, 61, 65, 67, 69, 75, 77, 79, 81, 83,

87, 89, 91, 93, 97, 101, 111

Changed, 53

Definition, 25

Integration, 43, 45

Partitioning, 81

Performance, 45

Segregation, 41

SOUP, 27, 33, 39

Software Of Unknown Provenance See SOUP, 27

SOFTWARE  PRODUCT, 19, 21, 23, 25, 27, 31,   49,

51, 53, 59, 61, 65, 73, 77, 85, 89, 91,  97

Definition, 25

Released, 51, 53

SOFTWARE  SYSTEM,  21,  25,  29, 31, 33, 37,  43,

53, 59, 61, 69, 73, 77, 79, 81, 83, 85, 89, 93,

95, 111

Definition, 25

Integration, 43 Requirements, 35 Testing, 45, 47

SOFTWARE UNIT, 25, 41, 43, 73, 77, 89,  91

Definition, 27

Integration, 43

Verification, 43

SOFTWARE UNIT Verification, 41

SOUP, 33, 35, 39, 41, 51, 55, 59, 75,  85

Change, 59

Configuration item, 31 Definition, 27

Designator, 59

Software item, 33

SYSTEM, 11,  19,  21,  23,  25,  31, 37, 39, 65, 73,

75, 79, 83, 85, 87, 101, 133

Configuration, 61

Definition, 27

Development plan, 31 Existing, 51

Released, 53

Requirements, 33, 35, 39, 41

TASK, 15, 17, 19, 23, 25, 29, 31, 73,    83, 93, 95,

97, 143

Completion of, 49

Configuration management, 35 Definition, 27

Deliverable, 19

Design and maintenance, 11 Maintenance, 51

Mapping, 15 Required, 15

Requirements, 17 Risk management, 33 Verification, 33

TRACEABILITY, 31, 57, 85, 87

Definition, 27

Verification, 25, 33, 35, 41, 43, 47, 49, 57, 61,

63, 69, 73, 75, 87, 91, 93, 97, 101, 113, 133,

145, 147

Definition, 27

VERSION, 49, 55, 59, 65, 95, 101

Definition, 27

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

__________

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304

Edition 1.1	2015-06

FINAL VERSION

colour inside

Medical device software – Software life cycle processes

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 2 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

## CONTENTS

FOREWORD	4

INTRODUCTION	6

INTRODUCTION to Amendment 1	8

- Scope	9

  - * Purpose	9

  - * Field of application	9

  - Relationship to other standards	9

  - Compliance	9

- * Normative references	10

- * Terms and definitions	10

- * General requirements	15

  - * Quality management system	15

  - * RISK MANAGEMENT	15

  - * Software safety classification	16

  - * LEGACY SOFTWARE	17

- Software development PROCESS	18

  - * Software development planning	18

  - * Software requirements analysis	21

  - * Software ARCHITECTURAL design	22

  - * Software detailed design	23

  - * SOFTWARE UNIT implementation	24

  - * Software integration and integration testing	25

  - * SOFTWARE SYSTEM testing	26

  - * Software release	27

- Software maintenance PROCESS	28

  - * Establish software maintenance plan	28

  - * Problem and modification analysis	29

  - * Modification implementation	30

- * Software RISK MANAGEMENT PROCESS	30

  - * Analysis of software contributing to  hazardous situations	30

  - RISK CONTROL measures	31

  - VERIFICATION of RISK CONTROL measures	31

  - RISK MANAGEMENT of software changes	31

- * Software configuration management PROCESS	32

  - * Configuration identification	32

  - * Change control	32

  - * Configuration status accounting	33

- * Software problem resolution PROCESS	33

  - Prepare PROBLEM REPORTS	33

  - Investigate the problem	33

  - Advise relevant parties	33

  - Use change control process	34

  - Maintain records	34

  - Analyse problems for trends	34

  - Verify software problem resolution	34

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 3 –

+AMD1:2015 CSV  IEC  2015

  - Test documentation contents	34

Annex A (informative)  Rationale for the requirements of  this standard	35

Annex B (informative)  Guidance on the provisions of  this standard	38

Annex C (informative)  Relationship to other standards	56

Annex D (informative)  Implementation	75

Bibliography	77

Index of defined terms	79

Figure 1 – Overview of software development PROCESSES  and ACTIVITIES	7

Figure 2 – Overview of software maintenance PROCESSES   and ACTIVITIES	7

Figure 3 – Assigning software safety classification	16

Figure B.2 – Pictorial representation of the relationship of HAZARD, sequence   of

events, HAZARDOUS SITUATION, and HARM – from ISO 14971:2007  Annex E	42

Figure B.1 – Example of partitioning of  SOFTWARE ITEMS	44

Figure C.1 – Relationship of key MEDICAL DEVICE standards to   IEC 62304	56

Figure C.2 – Software as part of  the V-model	59

Figure C.3 – Application of IEC 62304 with  IEC 61010-1	67

Table A.1 – Summary of requirements by software  safety class	37

Table B.1 – Development (model) strategies as defined in  ISO/IEC 12207	39

Table C.1 – Relationship to ISO 13485:2003	57

Table C.2 – Relationship to ISO 14971:2007	58

Table C.3 – Relationship to IEC 60601-1	61

Table C.5 – Relationship to ISO/IEC 12207	69

Table D.1 – Checklist for small companies without a  certified QMS	76

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 4 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

## INTERNATIONAL ELECTROTECHNICAL COMMISSION

MEDICAL DEVICE SOFTWARE –

SOFTWARE LIFE CYCLE PROCESSES

FOREWORD

- The International Electrotechnical  Commission (IEC) is  a  worldwide organization for standardization comprising all national electrotechnical committees (IEC National Committees). The object  of IEC  is to  promote  international co-operation on  all questions concerning standardization in the electrical and electronic fields. To   this end and in addition to other activities, IEC publishes International Standards, Technical Specifications, Technical Reports, Publicly Available Specifications (PAS) and Guides (hereafter referred to as “IEC Publication(s)”). Their preparation  is  entrusted to  technical  committees;  any  IEC National Committee interested in the subject dealt with may participate in this preparatory work. International, governmental and non-  governmental organizations liaising with  the  IEC also participate in  this  preparation.  IEC collaborates closely  with the International Organization for Standardization (ISO) in accordance with conditions determined by agreement between the two  organizations.

- The formal decisions or agreements of IEC on technical matters express, as nearly as possible, an international consensus of opinion on the relevant subjects since each technical committee has representation from all interested IEC National  Committees.

- IEC Publications have the form of recommendations for international use and are accepted by IEC National Committees in that sense. While all reasonable efforts are made to ensure that the technical content of IEC Publications is accurate, IEC cannot be held responsible for the way in which they are used or for any misinterpretation by any end  user.

- In order to promote international uniformity, IEC National Committees undertake to apply IEC Publications transparently to the maximum extent possible in their national  and  regional  publications.  Any  divergence  between any IEC Publication and the corresponding national or regional publication shall be clearly indicated in    the latter.

- IEC itself does not provide any attestation of conformity. Independent certification bodies provide conformity assessment services and, in some areas, access to IEC marks of conformity. IEC is not responsible for  any services carried out by independent certification   bodies.

- All users should ensure that they have the latest edition of this publication.

- No liability shall attach to IEC or its directors, employees, servants or agents including individual experts and members of its technical committees and IEC National Committees for any personal injury, property damage or other damage of any nature whatsoever, whether direct or indirect, or for costs (including  legal  fees)  and  expenses arising out of the publication, use of, or reliance upon, this IEC Publication or any  other  IEC  Publications.

- Attention is drawn to the Normative references cited in this publication. Use of the referenced publications is indispensable for the correct application of this   publication.

- Attention is drawn to the possibility that some of the elements of  this  IEC Publication may be  the  subject  of  patent rights. IEC shall not be held responsible for identifying any or all such patent rights.

DISCLAIMER

This Consolidated version is not an official IEC Standard and has been prepared for user convenience. Only the current versions of the standard and its amendment(s) are to be considered the official documents.

This Consolidated version of IEC 62304 bears the edition number 1.1. It consists of the first edition (2006-05) [documents 62A/523/FDIS and 62A/528/RVD] and its amendment 1 (2015-06) [documents 62A/1007/FDIS and 62A/1014/RVD]. The technical content is identical to the base edition and its amendment.

This Final version does not show where the technical content  is  modified  by amendment 1. A separate Redline version with  all  changes highlighted  is  available in this publication.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 5 –

+AMD1:2015 CSV  IEC 2015

International Standard IEC 62304 has been prepared by a joint working group of subcommittee 62A: Common aspects of electrical equipment used in medical practice, of IEC technical committee 62: Electrical equipment in medical practice and ISO Technical Committee 210, Quality management and corresponding general aspects for MEDICAL DEVICES. Table C.5 was prepared by ISO/IEC JTC 1/SC 7, Software and system  engineering.

It is published as a dual logo  standard.

This publication has been drafted in accordance with the ISO/IEC Directives, Part 2. In this standard the following print types are  used:

- requirements and definitions: in roman type;

- informative material appearing outside of tables, such as notes, examples and references:  in smaller type. Normative text of tables is also in a smaller  type;

- terms used throughout this standard that have been defined in Clause 3 and also given in   the index: in small capitals.

An asterisk (*) as the first character of a title or at the beginning of a paragraph indicates that there is guidance related to that item in Annex  B.

The committee has decided that the contents of the base publication and its amendment will remain unchanged until the stability date indicated on the IEC web site under "http://webstore.iec.ch" in the data related to the specific publication. At this  date,  the publication will be

- reconfirmed,

- withdrawn,

- replaced by a revised edition, or

- amended.

NOTE The attention of National Committees is drawn to the fact that equipment MANUFACTURERS and testing organizations  may  need  a   transitional  period  following  publication  of   a   new,   amended  or   revised  IEC  or   ISO publication in which to make products in accordance with the new requirements and to equip themselves for conducting new or revised tests. It is the recommendation of the committee that the content of this publication be adopted for mandatory implementation nationally not earlier than 3 years from the date of publication.

IMPORTANT – The 'colour inside' logo on the cover page of  this publication indicates that it contains colours which are considered to be useful for the correct understanding of its contents. Users should therefore print this document using a colour  printer.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 6 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

## INTRODUCTION

Software is often an integral part of MEDICAL DEVICE technology. Establishing the SAFETY and effectiveness of a MEDICAL DEVICE containing software requires knowledge of what the software  is intended to do and demonstration that the use of the software fulfils those intentions without causing any unacceptable RISKS.

This standard provides a framework of life cycle PROCESSES with ACTIVITIES and  TASKS  necessary for the safe design and maintenance of MEDICAL DEVICE SOFTWARE. This standard provides requirements for each life cycle PROCESS. Each life cycle PROCESS consists of a set of ACTIVITIES, with most ACTIVITIES consisting of a set of  TASKS.

As a basic foundation it is assumed that MEDICAL DEVICE SOFTWARE  is  developed  and  maintained within a quality management system (see 4.1) and a RISK MANAGEMENT system (see 4.2). The RISK MANAGEMENT PROCESS is already very well addressed by the  International Standard ISO 14971. Therefore IEC 62304 makes use of this advantage simply by a normative reference to ISO 14971. Some minor additional RISK MANAGEMENT requirements are needed for software, especially in the area of identification of contributing software factors related to HAZARDS. These requirements are summarized and captured in Clause 7 as the software RISK MANAGEMENT PROCESS.

Whether software is  a  contributing factor to  a HAZARDOUS  SITUATION  is determined during   the

HAZARD  identification  ACTIVITY  of  the  RISK  MANAGEMENT  PROCESS. HAZARDOUS  SITUATIONS that

could be indirectly caused by software (for example, by providing misleading information that could cause inappropriate treatment to be administered) need to be considered when  determining whether software is a contributing factor. The decision to use software to  control RISK is made during the RISK CONTROL ACTIVITY of the RISK MANAGEMENT PROCESS. The software RISK MANAGEMENT PROCESS required in this standard has to be embedded in the device RISK MANAGEMENT PROCESS according to ISO  14971.

The software development PROCESS consists of a number of ACTIVITIES. These ACTIVITIES are shown in Figure 1 and described in Clause 5. Because many incidents in the field are related to service or maintenance of MEDICAL DEVICE SYSTEMS  including inappropriate software updates  and upgrades, the software maintenance PROCESS is considered to be as important as the software development PROCESS. The software maintenance PROCESS is very similar to the software development PROCESS. It is shown in Figure 2 and described in Clause   6.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 7 –

+AMD1:2015 CSV  IEC 2015

SYSTEM development ACTIVITIES (including RISK  MANAGEMENT)

5.1

Software development planning

5.2

Software requirements analysis

5.3

Software

ARCHITECTURAL

design

5.4

Software detailed design

5.5

Software UNIT implementation and VERIFICATION

5.6

Software integration and integration testing

5.7

Software SYSTEM

testing

5.8

Software release

Activities outside the scope of this standard

Customer needs

Customer needs satisfied

7  Software RISK MANAGEMENT

8  Software configuration management

9  Software problem resolution

IEC  722/06

Figure 1 – Overview of software development PROCESSES and  ACTIVITIES

Maintenance request

System maintenance ACTIVITIES (including RISK MANAGEMENT)

6.1

Establish software maintenance

plan

6.2

Problem and modification analysis

5.3

Software

ARCHITECTURAL

design

5.4

Software detailed design

5.5

Software UNIT

implementation and

VERIFICATION

5.6

Software integration and integration testing

5.7

Software SYSTEM

testing

5.8

Software release

Activities outside the scope of this standard

Request satisfied

7 Software RISK MANAGEMENT

6.3 Modification implementation

8 Software configuration management

9 Software problem resolution

IEC  723/06

Figure 2 – Overview of software maintenance PROCESSES and  ACTIVITIES

This  standard  identifies  two  additional  PROCESSES  considered  essential  for  developing safe

MEDICAL DEVICE SOFTWARE. They are the software configuration management PROCESS     (Clause

8) and the software problem resolution PROCESS (Clause  9).

Amendment 1 updates the standard to add requirements to deal with LEGACY SOFTWARE, where the software design is prior to the existence of the current version, to assist manufacturers who must  show  compliance  to   the   standard  to   meet  European  Directives.  Software      safety

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 8 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

classification changes include clarification of requirements and updating of the software safety classification to include a risk-based approach.

This standard does not specify an organizational structure for the MANUFACTURER  or which part   of the organization is to perform which PROCESS, ACTIVITY, or TASK. This standard requires only that the PROCESS, ACTIVITY, or TASK be completed to establish compliance with this   standard.

This standard does not prescribe the name, format, or explicit content of the documentation to   be produced. This standard requires documentation of TASKS, but the decision of how  to  package this documentation is left to the user of the  standard.

This standard does not prescribe a specific life cycle model. The users of this standard are responsible for selecting a life cycle model for the software project and for mapping the PROCESSES, ACTIVITIES, and TASKS in this standard onto that  model.

Annex A provides rationale for the clauses of this standard. Annex B provides guidance on the provisions of this standard.

For the purposes of this standard:

- “shall” means that compliance with a requirement is mandatory for compliance with this standard;

- “should” means that compliance with a requirement is recommended but is not mandatory for compliance with this standard;

- “may” is used to describe a permissible way to achieve compliance with a   requirement;

- “establish” means to define, document, and implement; and

- where this standard uses the term “as appropriate” in conjunction with a required PROCESS, ACTIVITY, TASK or output, the intention is that the MANUFACTURER shall use the PROCESS, ACTIVITY, TASK or output unless the MANUFACTURER can document a justification for not so doing.

## INTRODUCTION to Amendment 1

The first edition of IEC 62304 was published in 2006. This amendment is intended to add requirements to deal with LEGACY SOFTWARE, where the software design is  prior  to  the existence of the current version, to assist manufacturers who must show compliance to the standard to meet European Directives. Software safety classification changes needed for this amendment include clarification of requirements and updating of the software safety classification to include a risk-based approach. Work is continuing in parallel to develop the second edition of IEC 62304.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 9 –

+AMD1:2015 CSV  IEC 2015

# MEDICAL DEVICE SOFTWARE – SOFTWARE LIFE CYCLE PROCESSES

### Scope

#### * Purpose

This standard defines the life cycle requirements for MEDICAL DEVICE SOFTWARE. The set of PROCESSES, ACTIVITIES, and TASKS described in this standard establishes a common framework for MEDICAL DEVICE SOFTWARE life cycle  PROCESSES.

#### * Field of application

This standard applies to the development and maintenance of MEDICAL DEVICE SOFTWARE when software is itself a MEDICAL DEVICE or when software is an embedded or integral part of the final MEDICAL DEVICE.

NOTE 1 This standard can be used in the development  and  maintenance  of  software  that  is  itself  a  medical device.   However,  additional development activities are needed at the system level before this type of software can    be placed into service.  These system activities are not covered by this standard, but can be found in IEC 82304-11  [22].

This standard describes PROCESSES that are intended to be applied to software which executes on a processor or which is executed by other software (for example an interpreter) which executes on a processor.

This standard applies regardless of the persistent storage device(s) used to store the software (for example: hard disk, optical disk, permanent or flash  memory).

This standard applies regardless of the method of delivery of the software (for example: transmission by network or email, optical disk, flash memory or EEPROM). The method of software delivery itself is not considered MEDICAL DEVICE   SOFTWARE.

This standard does not cover validation and final release of the MEDICAL DEVICE, even when the

MEDICAL DEVICE consists entirely of  software.

NOTE 2 If a medical device incorporates embedded software intended to be executed on a processor,  the  requirements of this standard apply to the software, including the requirements concerning software of unknown provenance (see 8.1.2).

NOTE 3 Validation and other development activities are needed at the  system  level  before  the  software  and  medical device can be placed into service.   These system activities are not covered by this standard, but can be   found in related product standards (e.g., IEC 60601-1, IEC 82304-1,    etc.).

#### Relationship to other standards

This MEDICAL DEVICE SOFTWARE life cycle standard is to be used together with other appropriate standards when developing a MEDICAL DEVICE. Annex C shows the relationship between this standard and other relevant standards.

#### Compliance

Compliance with this standard is defined as implementing all of the PROCESSES, ACTIVITIES,   and

TASKS identified in this standard in accordance with the software safety  class.

NOTE The software safety classes assigned to each requirement are identified in the normative text following the requirement.

1    In preparation.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 10 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Compliance is determined by inspection of all documentation required by  this  standard  including the RISK MANAGEMENT FILE, and assessment of the PROCESSES, ACTIVITIES and TASKS required for the software safety class.

NOTE 1    This assessment could be carried out by internal or external    audit.

NOTE 2    Although the specified PROCESSES, ACTIVITIES, and TASKS  are performed, flexibility exists in the methods     of implementing these PROCESSES and performing these ACTIVITIES and   TASKS.

NOTE 3 Where any requirements contain “as appropriate” and were not performed,  documentation  for  the  justification is necessary for this  assessment.

NOTE 4 The term “conformance” is used in ISO/IEC 12207 where the term “compliance” is used in this standard. NOTE 5    For compliance of LEGACY SOFTWARE see  4.4.

### * Normative references

The following referenced documents are indispensable for the application of this document. For dated references, only the edition cited applies. For  undated references, the latest edition of   the referenced document (including any amendments) applies.

ISO 14971, Medical devices – Application of risk management to medical  devices.

### * Terms and definitions

For the purposes of this document, the following terms and definitions  apply.

#### 3.1

ACTIVITY

a set of one or more interrelated or interacting  TASKS

#### 3.2

ANOMALY

any condition that deviates from the expected based on requirements specifications, design documents, standards, etc. or from someone’s perceptions or experiences. ANOMALIES may be found during, but not limited to, the review, test, analysis,  compilation,  or  use  of  MEDICAL DEVICE SOFTWARE or applicable  documentation

NOTE   Based on IEEE 1044:1993, definition 3.1.

#### 3.3

ARCHITECTURE

organizational structure of a SYSTEM or component [IEEE 610.12:1990]

#### 3.4

CHANGE REQUEST

a documented specification of a change to be made to a MEDICAL DEVICE    SOFTWARE

#### 3.5

CONFIGURATION ITEM

entity that can be uniquely identified at a given reference  point

NOTE    Based on ISO/IEC 12207:2008, 4,7.

#### 3.6

DELIVERABLE

required result or output (includes documentation) of an ACTIVITY or  TASK

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 11 –

+AMD1:2015 CSV  IEC 2015

#### 3.7

EVALUATION

a systematic determination of the extent to which an entity meets its specified criteria [ISO/IEC 12207:2008, 4.12]

#### 3.8

HARM

physical injury, damage, or both to the health of people or damage to property or the  environment

[ISO 14971:2007, 2.2]

#### 3.9

HAZARD

potential source of HARM

[ISO 14971:2007, 2.3]

#### 3.10

MANUFACTURER

natural or legal person with responsibility for designing, manufacturing, packaging, or labelling    a MEDICAL DEVICE; assembling a SYSTEM; or adapting a MEDICAL  DEVICE  before it is  placed on  the market and/or put into service, regardless  of  whether  these operations are carried out by that person or by a third party on that person’s  behalf

NOTE 1 Attention is drawn to the fact that the provisions of national  or  regional  regulations  can  apply  to  the definition of manufacturer.

NOTE 2    For a definition of labelling, see ISO 13485:2003, definition   3.6.

[ISO 14971:2007, 2,8]

#### 3.11

MEDICAL DEVICE

any instrument, apparatus, implement, machine, appliance, implant, in vitro reagent or calibrator, software, material or other similar or related article, intended by the MANUFACTURER    to be used, alone or in combination, for human beings for one or more  of  the  specific purpose(s) of

- diagnosis, prevention, monitoring, treatment or alleviation of  disease,

- diagnosis, monitoring, treatment, alleviation of or compensation for an  injury,

- investigation, replacement, modification, or  support of  the anatomy or  of  a     physiological

PROCESS,

- supporting or sustaining life,

- control of conception,

- disinfection of MEDICAL DEVICES,

- providing information for medical purposes by means of in vitro examination of specimens derived from the human body,

and which does not achieve its primary intended action in or on the human body by pharmacological, immunological or metabolic means, but which may be assisted in its function by such means

NOTE 1 This definition has been developed by the Global Harmonization Task Force (GHTF). See bibliographic reference [15] (in ISO  13485:2003).

[ISO 13485:2003, definition 3.7]

NOTE 2        Some differences can occur in the definitions used in regulations of each country.

NOTE  3	In  conjunction  with  IEC 60601-1:2005  and  IEC 60601-1:2005/AMD1:2012  the  term     “medical  device” assumes the same meaning as ME EQUIPMENT or ME SYSTEM (which are defined terms of IEC 60601-1).

#### 3.12

MEDICAL DEVICE SOFTWARE

– 12 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

SOFTWARE  SYSTEM  that  has  been  developed  for  the  purpose  of  being  incorporated  into the

MEDICAL DEVICE being developed or that is intended for use as a   MEDICAL DEVICE

NOTE    This includes a MEDICAL DEVICE  software product, which then is a MEDICAL DEVICE  in its own   right.

#### 3.13

PROBLEM REPORT

a record of actual or potential behaviour of a MEDICAL DEVICE SOFTWARE that a user or other interested person believes to be unsafe, inappropriate for the intended use or contrary to specification

NOTE 1 This standard does not require that every PROBLEM REPORT results in a change to the MEDICAL DEVICE SOFTWARE. A MANUFACTURER can reject a PROBLEM REPORT as a misunderstanding, error or insignificant event.

NOTE 2      A  PROBLEM REPORT  can relate to a  released MEDICAL DEVICE  SOFTWARE  or to a MEDICAL DEVICE  SOFTWARE

that is still under  development.

NOTE 3  This standard requires the MANUFACTURER  to perform extra decision making steps (see Clause 6) for a

PROBLEM REPORT relating to a released product to ensure that regulatory actions are identified and implemented.

#### 3.14

PROCESS

a set of interrelated or interacting ACTIVITIES that transform inputs into outputs [ISO 9000:2000, definition 3.4.1]

NOTE    The term “ACTIVITIES” covers use of  resources.

#### 3.15

REGRESSION TESTING

the testing required to determine that a change to a SYSTEM component has not adversely affected functionality, reliability or performance and has not introduced additional  defects

[ISO/IEC 90003:2004, definition 3.11]

#### 3.16

RISK

combination of the probability of occurrence of HARM and the severity of that   HARM

[ISO 14971:2007, 2.16]

#### 3.17

RISK ANALYSIS

systematic use of available information to identify HAZARDS and to estimate the   RISK

[ISO 14971:2007, 2.17]

#### 3.18

RISK CONTROL

PROCESS in which decisions are made and RISKS are reduced to, or maintained within, specified levels

[ISO 14971:2007, 2.19]

#### 3.19

RISK MANAGEMENT

systematic application of management policies, procedures, and practices to the TASKS of analyzing, evaluating, and controlling RISK

[ISO 14971:2007,  2.22, modified – The phrase "and monitoring" has been  removed]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 13 –

+AMD1:2015 CSV  IEC 2015

#### 3.20

RISK MANAGEMENT FILE

set of records and other documents, not necessarily contiguous, that are produced by a RISK MANAGEMENT PROCESS

[ISO 14971:2007, 2.23]

#### 3.21

SAFETY

freedom from unacceptable RISK

[ISO 14971:2007, 2.24]

#### 3.22

SECURITY

protection of information and data so that unauthorized persons or systems cannot read or modify them an authorized persons or systems are not denied access to  them

NOTE  Based on ISO/IEC 12207: 2008,  4.39.

#### 3.23

SERIOUS INJURY

injury or illness that:

- is life threatening,

- results in permanent impairment of a body function or permanent damage to a body structure, or

- necessitates medical or surgical intervention to prevent permanent impairment of a body function or permanent damage to a body structure

NOTE Permanent impairment means an irreversible impairment or damage  to  a  body  structure  or  function  excluding trivial impairment or  damage.

#### 3.24

SOFTWARE DEVELOPMENT LIFE CYCLE  MODEL

conceptual structure spanning the life of the software from definition of its requirements to its release, which:

- identifies the PROCESS, ACTIVITIES and TASKS involved in development of MEDICAL DEVICE SOFTWARE,

- describes the sequence of and dependency between ACTIVITIES and TASKS,  and

- identifies the milestones at which the completeness of specified DELIVERABLES is   verified.

NOTE    Based on ISO/IEC 12207:1995, definition  3.11

#### 3.25

SOFTWARE ITEM

any identifiable part of a computer program, i.e., source code,  object  code,  control  code, control data, or a collection of these  items

NOTE  Three  terms  identify the  software  decomposition. The top level is  the SOFTWARE SYSTEM. The lowest level  that is not further decomposed is the SOFTWARE UNIT. All levels of composition, including the top and bottom levels,   can be called SOFTWARE ITEMS. A SOFTWARE SYSTEM, then, is composed of one or more SOFTWARE ITEMS, and each SOFTWARE ITEM  is  composed of  one or more SOFTWARE UNITS or decomposable SOFTWARE ITEMS. The responsibility   is left to the MANUFACTURER to provide the granularity of the SOFTWARE ITEMS and SOFTWARE UNITS.

NOTE 2   Based on ISO/IEC 90003:2004, 3.14 and ISO/IEC 12207:2008,   4.41

#### 3.26

Not used

#### 3.27

SOFTWARE SYSTEM

integrated collection of SOFTWARE ITEMS organized to accomplish a specific function or set of functions

#### 3.28

SOFTWARE UNIT

– 14 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

SOFTWARE ITEM that is not subdivided into other  items

NOTE    The granularity of SOFTWARE UNITS  is defined by the MANUFACTURER (see  B.3).

#### 3.29

SOUP

#### software of unknown provenance (acronym)

SOFTWARE ITEM that is already developed and generally available and that has not been developed for the purpose of being incorporated into the MEDICAL DEVICE (also known as “off- the-shelf software”) or SOFTWARE ITEM previously developed for which adequate records of the development PROCESSES are not  available

NOTE    A MEDICAL DEVICE SOFTWARE SYSTEM in itself cannot be claimed to be   SOUP.

#### 3.30

SYSTEM

integrated composite consisting of one or more of the PROCESSES, hardware,  software,  facilities, and people, that provides a capability to satisfy a stated need or  objective

NOTE   Based on ISO/IEC ISO/IEC 12207:2008,  4.48.

#### 3.31

TASK

a single piece of work that needs to be  done

#### 3.32

TRACEABILITY

degree to which a relationship can be established between two or more products of the development PROCESS

[IEEE 610.12:1990]

NOTE   Requirements, architecture, risk  control measures, etc.  are  examples  of  deliverables  of  the    development

PROCESS.

#### 3.33

VERIFICATION

confirmation through provision of objective evidence that specified requirements have been fulfilled

NOTE 1    “Verified” is used to designate the corresponding   status.

[ISO 9000:2000, definition 3.8.4]

NOTE 2    In  design  and  development,  VERIFICATION  concerns  the  PROCESS  of  examining  the  result  of  a      given

ACTIVITY to determine conformity with the stated requirement for that   ACTIVITY.

#### 3.34

VERSION

identified instance of a CONFIGURATION  ITEM

NOTE 1 Modification to a VERSION of MEDICAL DEVICE SOFTWARE, resulting in a new VERSION, requires software configuration management action.

NOTE 2    Based on ISO/IEC 12207:2008, 4.56.

#### 3.35

HAZARDOUS SITUATION

circumstance  in  which  people,  property  or  the  environment  are  exposed  to  one  or    more

HAZARD(S)

[SOURCE: ISO 14971:2007, 2.4]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 15 –

+AMD1:2015 CSV  IEC 2015

#### 3.36

LEGACY SOFTWARE

MEDICAL DEVICE SOFTWARE which was  legally placed on the market and is  still marketed today  but for which there is insufficient objective evidence that it was  developed in  compliance with  the current version of this standard

#### 3.37

RELEASE

particular VERSION of a CONFIGURATION ITEM that is made available for a specific   purpose

NOTE    Based on ISO/IEC 12207:2008, definition  4.35.

#### 3.38

RESIDUAL RISK

RISK remaining after RISK CONTROL measures have been  taken

NOTE 1    Adapted from ISO/IEC Guide 51:1999, definition  3.9.

NOTE 2 ISO/IEC Guide 51:1999, definition 3.9 uses the term “protective measures” rather than “RISK CONTROL measures.” However, in the context of this International Standard, “protective measures” are only one option for controlling RISK  as described in 6.2 [of ISO  14971:2007].

[SOURCE: ISO 14971:2007, 2.15].

#### 3.39

RISK ESTIMATION

PROCESS  used to assign values to the probability of occurrence of HARM and the severity of    that

HARM

[SOURCE: ISO 14971:2007 2.20]

#### 3.40

RISK EVALUATION

PROCESS of comparing the estimated RISK against given RISK criteria to determine the acceptability of the RISK

[SOURCE: ISO 14971:2007 2.21]

### * General requirements

#### * Quality management system

The MANUFACTURER of MEDICAL DEVICE SOFTWARE shall demonstrate the ability  to  provide MEDICAL DEVICE SOFTWARE that consistently meets customer requirements and applicable regulatory requirements.

NOTE 1       Demonstration of this ability can be by the use of a quality management system that complies with:

- ISO 13485 [8];  or

- a national quality management system standard;   or

- a quality management system required by national   regulation.

NOTE 2 Guidance for applying quality management system requirements to  software  can  be  found  in  ISO/IEC 90003 [15].

  - * RISK MANAGEMENT

The MANUFACTURER shall apply a RISK MANAGEMENT PROCESS complying with ISO   14971.

#### * Software safety classification

– 16 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

- The MANUFACTURER shall assign to each SOFTWARE SYSTEM a software safety class (A, B,    or

C) according to the RISK of HARM to the patient, operator, or other people resulting from a HAZARDOUS SITUATION to which the SOFTWARE SYSTEM can contribute in a  worst-case-  scenario as indicated in Figure 3.

IEC

#### Figure 3 – Assigning software safety classification

The SOFTWARE SYSTEM is software safety class A  if:

- the SOFTWARE SYSTEM cannot contribute to a HAZARDOUS SITUATION;   or

- the SOFTWARE SYSTEM can contribute to a HAZARDOUS SITUATION which does not result in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM.

The SOFTWARE SYSTEM is software safety class B  if:

- the SOFTWARE SYSTEM can contribute to a HAZARDOUS SITUATION which  results  in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM and the resulting possible HARM is non-SERIOUS   INJURY.

The SOFTWARE SYSTEM is software safety class C  if:

- the SOFTWARE SYSTEM can contribute to a HAZARDOUS SITUATION which  results  in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM and the resulting possible HARM is death or SERIOUS   INJURY.

For a SOFTWARE SYSTEM initially classified as software safety class B or C, the MANUFACTURER may implement additional RISK CONTROL measures external to the SOFTWARE SYSTEM (including revising the system architecture containing the SOFTWARE SYSTEM) and subsequently assign a new software safety classification to the SOFTWARE  SYSTEM.

NOTE 1 External RISK CONTROL measures can be hardware, an independent SOFTWARE SYSTEM, health care  procedures, or other means to minimize that software can contribute to a HAZARDOUS SITUATION.

NOTE 2        See ISO 14971:2007 subclause 3.2, Management Responsibilities, for the definition of risk acceptability.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 17 –

+AMD1:2015 CSV  IEC 2015

- Not used.

- The MANUFACTURER shall document the software safety class assigned to each SOFTWARE SYSTEM in the RISK MANAGEMENT  FILE.

- When a SOFTWARE SYSTEM is  decomposed into SOFTWARE  ITEMS, and when a  SOFTWARE  ITEM is decomposed into further SOFTWARE ITEMS, such SOFTWARE ITEMS shall inherit the software safety classification of the original SOFTWARE ITEM (or SOFTWARE  SYSTEM) unless  the MANUFACTURER documents a rationale for classification into a different software safety class (software safety classes assigned according to 4.3 a) replacing “SOFTWARE SYSTEM” with “SOFTWARE ITEM”). Such a rationale shall explain how the new SOFTWARE ITEMS are segregated so that they may be classified separately.

- The MANUFACTURER shall document the software safety class of each SOFTWARE ITEM if that class is different from the class of the SOFTWARE ITEM from which it was created by decomposition.

- For compliance with this standard, when applying this standard to a group of  SOFTWARE ITEMS, the MANUFACTURER shall use the PROCESSES and TASKS which are required by the classification of the highest-classified SOFTWARE ITEM in the group unless  the MANUFACTURER documents in the RISK MANAGEMENT FILE a rationale for using a lower classification.

- For each SOFTWARE SYSTEM, until a software safety class is assigned,  Class  C requirements shall apply.

NOTE In the clauses and subclauses that follow, the  software  safety  classesfor  which  a  specific  requirement applies are identified following the requirement in the form [Class . . .].

  - * LEGACY SOFTWARE

#### General

As an alternative to applying Clauses 5 through 9 of this standard, compliance of LEGACY SOFTWARE may be demonstrated as indicated in 4.4.2 to  4.4.5.

    - RISK MANAGEMENT ACTIVITIES

In accordance with 4.2 of this standard, the MANUFACTURER   shall:

- assess any feedback, including post-production information, on LEGACY SOFTWARE regarding incidents and / or near incidents, both from inside its own organization and / or from users;

- perform RISK MANAGEMENT ACTIVITIES associated with continued use of the  LEGACY  SOFTWARE, considering the following aspects:

  - integration of the LEGACY SOFTWARE in the overall MEDICAL DEVICE   architecture;

  - continuing validity of RISK CONTROL measures, implemented as part of the LEGACY SOFTWARE;

  - identification of HAZARDOUS SITUATIONS associated with the continued use of the LEGACY SOFTWARE;

  - identification of potential causes of the LEGACY SOFTWARE contributing to a HAZARDOUS SITUATION;

  - definition of  RISK  CONTROL  measures for each potential cause of  the LEGACY   SOFTWARE

contributing to a HAZARDOUS  SITUATION.

#### Gap analysis

Based on the software safety class of the LEGACY SOFTWARE (see 4.3), the MANUFACTURER shall perform a gap analysis of available DELIVERABLES against those required according to 5.2, 5.3, 5.7, and Clause 7.

- The MANUFACTURER shall assess the continuing validity of available  DELIVERABLES.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 18 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

- Where  gaps  are  identified,  the  MANUFACTURER  shall  EVALUATE  the  potential  reduction in

RISK resulting from the generation of the missing DELIVERABLES and associated   ACTIVITIES.

- Based on this evaluation, the MANUFACTURER shall determine the DELIVERABLES  to  be  created and associated ACTIVITIES to be performed. The minimum DELIVERABLE shall be SOFTWARE SYSTEM test records (see  5.7.5).

NOTE Such gap analysis should assure that RISK CONTROL measures, implemented  in  LEGACY  SOFTWARE,  are included in the software  requirements.

#### Gap closure activities

- The MANUFACTURER shall establish and execute a plan to generate the identified DELIVERABLES. Where available, objective evidence may be used to generate required DELIVERABLES without performing ACTIVITIES required by 5.2, 5.3, 5.7 and Clause   7.

NOTE        A plan on how to address the identified gaps can be included in a software maintenance plan (see 6.1).

- The plan shall address the use of the problem resolution PROCESS for handling problems detected in the LEGACY SOFTWARE and DELIVERABLES in accordance with Clause   9.

- Changes to the LEGACY SOFTWARE shall be performed in accordance with Clause   6.

    - Rationale for use of LEGACY  SOFTWARE

The MANUFACTURER shall document the VERSION of the LEGACY SOFTWARE together with a  rationale for the continued use of the LEGACY SOFTWARE based on the outputs of   4.4.

NOTE    Fulfilling 4.4 enables further use of LEGACY SOFTWARE in accordance with IEC   62304.

- Software development PROCESS

#### * Software development planning

    - Software development plan

The MANUFACTURER shall establish a software development plan (or plans) for conducting the ACTIVITIES of the software development PROCESS appropriate to the scope, magnitude, and software safety classifications of the SOFTWARE SYSTEM to be developed. The sOFTWARE DEVELOPMENT LIFE CYCLE MODEL shall either be fully defined or be referenced in the plan (or plans). The plan shall address the following:

- the PROCESSES to be used in the development of the SOFTWARE SYSTEM    (see Note 4);

- the DELIVERABLES (includes documentation) of the ACTIVITIES and  TASKS;

- TRACEABILITY   between  SYSTEM   requirements,  software  requirements,  SOFTWARE  SYSTEM

test, and RISK CONTROL measures implemented in  software;

- software configuration and change management, including SOUP CONFIGURATION ITEMS and software used to support development; and

- software problem resolution for handling problems detected in the  MEDICAL  DEVICE SOFTWARE, DELIVERABLES and ACTIVITIES at each stage of the life   cycle.

[Class A, B, C]

NOTE 1 The SOFTWARE DEVELOPMENT LIFE CYCLE MODEL can identify different  elements  (PROCESSES,  ACTIVITIES, TASKS and DELIVERABLES) for different SOFTWARE ITEMS according to the software safety classification of each SOFTWARE ITEM of the SOFTWARE SYSTEM.

NOTE 2  These ACTIVITIES and TASKS  can overlap or interact and can be performed iteratively or recursively. It is not  the intent to imply that a specific life cycle model should be used.

NOTE 3 Other PROCESSES are described in this standard separately from the development PROCESS. This does not imply that they must be implemented as separate ACTIVITIES and TASKS. The ACTIVITIES and TASKS of the other PROCESSES  can be integrated into the development  PROCESS.

NOTE 4  The software development plan can reference existing PROCESSES  or    define new ones.

NOTE 5     The software development plan may be integrated in an overall SYSTEM  development plan.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 19 –

+AMD1:2015 CSV  IEC 2015

#### Keep software development plan updated

The MANUFACTURER  shall update the plan as development proceeds as appropriate. [Class A,   B, C]

#### Software development plan reference to SYSTEM design and  development

- As inputs for software development, SYSTEM requirements shall be referenced in the  software development plan by the  MANUFACTURER.

- In the software development plan, the MANUFACTURER shall include or reference procedures for coordinating the software development with the system development necessary  to  satisfy 4.1 (such as system integration, verification, and  validation).

[Class A, B, C]

NOTE   There might not  be  a  difference between SOFTWARE SYSTEM  requirements and  SYSTEM  requirements if    the

SOFTWARE SYSTEM  is a stand alone SYSTEM  (software-only device).

#### Software development standards, methods and tools planning

The MANUFACTURER shall include or reference in the software development  plan:

- standards,

- methods, and

- tools

associated with the development of SOFTWARE ITEMS of class C. [Class   C]

#### Software integration and integration testing planning

The MANUFACTURER shall include or reference in the software development plan, a plan to integrate the SOFTWARE ITEMS (including SOUP) and perform testing during integration. [Class B, C]

NOTE 1          It is acceptable to combine integration testing and SOFTWARE SYSTEM  testing into a single plan and set of

ACTIVITIES.

NOTE 2    See 5.6.

    - Software VERIFICATION planning

The MANUFACTURER shall include or reference in the software development plan the following

VERIFICATION information:

- DELIVERABLES  requiring VERIFICATION;

- the required VERIFICATION TASKS for each life cycle  ACTIVITY;

- milestones at which the DELIVERABLES are VERIFIED;  and

- the acceptance criteria for VERIFICATION of the DELIVERABLES. [Class A, B, C]

    - Software RISK MANAGEMENT  planning

The MANUFACTURER shall include or reference in the software development plan, a  plan  to conduct the ACTIVITIES and TASKS of the software RISK MANAGEMENT PROCESS, including the management of RISKS relating to SOUP. [Class A, B,  C]

NOTE    See Clause 7.

#### Documentation planning

The MANUFACTURER shall include or reference in the software development plan information about the documents to be produced during the software development life cycle. For each

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 20 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

identified document or type of document the following information shall be included or referenced:

- title, name or naming convention;

- purpose;

- procedures and responsibilities for development, review, approval and modification. [Class A, B, C]

NOTE    See Clause 8 for consideration of configuration management of   documentation.

#### Software configuration management planning

The MANUFACTURER  shall include or reference software configuration management information  in the software development plan. The software configuration management information shall include or reference:

- the classes, types, categories or lists of items to be  controlled;

- the software configuration management ACTIVITIES and  TASKS;

- the   organization(s)   responsible   for   performing  software   configuration    management

ACTIVITIES;

- their relationship with other organizations, such as software development or  maintenance;

- when the items are to be placed under configuration control;  and

- when the problem resolution PROCESS is to be used. [Class A, B, C]

NOTE    See Clause 8.

#### Supporting items to be controlled

The items to be controlled shall include tools, items or settings, used to develop the MEDICAL DEVICE SOFTWARE, which could impact the MEDICAL DEVICE SOFTWARE. [Class B,   C]

NOTE 1 Examples of such items include compiler/assembler versions, make files, batch files, and  specific  environment settings.

NOTE 2    See Clause 8.

    - Software CONFIGURATION ITEM  control  before VERIFICATION

The MANUFACTURER shall plan to place CONFIGURATION ITEMS under configuration management control before they are VERIFIED. [Class B,  C]

#### Identification and avoidance of common software defects

The MANUFACTURER  shall include or reference in the software development plan a  procedure  for:

- identifying categories of defects that may be introduced based on the  selected  programming technology that are relevant to their SOFTWARE SYSTEM;  and

- documenting evidence that demonstrates that these defects do not contribute to unacceptable RISK.

NOTE    See  Annex  B  of  IEC  TR  80002-1:2009  for  examples  of  categories  of  defects  or  causes  contributing to

HAZARDOUS SITUATIONS.

[Class B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 21 –

+AMD1:2015 CSV  IEC 2015

#### * Software requirements analysis

    - Define and document software requirements from SYSTEM  requirements

For each SOFTWARE SYSTEM of the MEDICAL DEVICE, the MANUFACTURER shall define and  document SOFTWARE SYSTEM requirements from the SYSTEM level requirements. [Class A, B,   C]

NOTE   There might not  be  a  difference between SOFTWARE SYSTEM  requirements and  SYSTEM  requirements if    the

SOFTWARE SYSTEM  is a stand alone SYSTEM  (software-only device).

#### Software requirements content

As appropriate to the MEDICAL DEVICE SOFTWARE, the MANUFACTURER shall include  in  the  software requirements:

- functional and capability requirements;

NOTE 1  Examples  include:

- performance (e.g., purpose of software, timing  requirements),

- physical characteristics (e.g., code language, platform, operating   system),

- computing environment (e.g., hardware, memory size, processing unit, time zone, network infrastructure) under which the software is to perform,  and

- need for compatibility with upgrades or multiple SOUP  or other device   versions.

- SOFTWARE SYSTEM inputs and  outputs;

NOTE 2  Examples  include:

- data characteristics (e.g., numerical, alpha-numeric,  format)

- ranges,

- limits, and

- defaults.

- interfaces between the SOFTWARE SYSTEM and other  SYSTEMS;

- software-driven alarms, warnings, and operator messages;

- SECURITY requirements;

NOTE 3  Examples  include:

- those related to the compromise of sensitive   information,

- authentication,

- authorization,

- audit trail, and

- communication integrity,

- system  security/malware protection.

- user interface requirements implemented by software;

NOTE 4  Examples include those related  to:

- support for manual  operations,

- human-equipment interactions,

- constraints on personnel,  and

- areas needing concentrated human  attention.

NOTE 5 Information regarding usability engineering requirements can be found in IEC 62366-1 [21] among others  (e.g., IEC 60601-1-6 [3]).

- data definition and database requirements;

NOTE 6  Examples  include:

- form;

- fit;

- function.

- installation and acceptance requirements of the delivered MEDICAL DEVICE SOFTWARE at the operation and maintenance site or sites;

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 22 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

- requirements related to methods of operation and  maintenance;

- requirements related to IT-network aspects;

NOTE 9    Examples include those related  to:

- networked alarms, warnings, and operator  messages;

- network protocols;

- handling of unavailability of network  services.

- user maintenance requirements; and

- regulatory requirements.

NOTE 10    The requirements in a) through l) can  overlap.

[Class A, B, C]

NOTE 7       All of these requirements might not be available at the beginning of the software development.

NOTE 8 Among others, ISO/IEC 25010 [12] provides information on quality characteristics that may be useful in defining software requirements.

    - Include RISK CONTROL measures in software  requirements

The MANUFACTURER shall include RISK CONTROL measures implemented in software in the requirements as appropriate to the MEDICAL DEVICE SOFTWARE. [Class B,   C]

NOTE  These  requirements  might  not  be  available at  the beginning of  the software development and can change  as the software is designed and RISK CONTROL  measures are further   defined.

    - Re-EVALUATE  MEDICAL DEVICE  RISK ANALYSIS

The MANUFACTURER shall re-EVALUATE the MEDICAL DEVICE RISK ANALYSIS when software requirements are established and update it as appropriate. [Class A, B,  C]

#### Update requirements

The MANUFACTURER shall ensure that  existing  requirements, including  SYSTEM  requirements, are re-EVALUATED and updated as appropriate as a result of the software requirements analysis ACTIVITY. [Class A, B, C]

#### Verify software requirements

The MANUFACTURER shall verify and document that the software  requirements:

- implement SYSTEM requirements including those relating to RISK   CONTROL;

- do not contradict one another;

- are expressed in terms that avoid ambiguity;

- are stated in terms that permit establishment of test criteria and performance of  tests;

- can be uniquely identified; and

- are traceable to SYSTEM requirements or other source. [Class A, B, C]

NOTE    This standard does not require the use of a formal specification   language.

  - * Software ARCHITECTURAL design

    - Transform software requirements into an ARCHITECTURE

The MANUFACTURER shall transform the requirements for the MEDICAL DEVICE SOFTWARE into a documented ARCHITECTURE that describes the software’s structure and identifies the SOFTWARE ITEMS. [Class B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 23 –

+AMD1:2015 CSV  IEC 2015

    - Develop an ARCHITECTURE for the interfaces of SOFTWARE   ITEMS

The MANUFACTURER shall develop and document an ARCHITECTURE  for the interfaces between  the SOFTWARE ITEMS and the components external to the SOFTWARE ITEMS (both software and hardware), and between the SOFTWARE ITEMS. [Class B,  C]

#### Specify functional and performance requirements of SOUP  item

If a SOFTWARE ITEM is identified as SOUP, the MANUFACTURER shall specify functional and performance requirements for  the  SOUP  item  that are necessary for its intended use. [Class    B, C]

    - Specify SYSTEM hardware and software required by SOUP   item

If a SOFTWARE ITEM is identified as SOUP, the MANUFACTURER shall specify the SYSTEM hardware and software necessary to support the proper operation of the SOUP item. [Class B,   C]

NOTE	Examples include processor type and speed, memory type and size, SYSTEM  software    type, communication and display software  requirements.

    - Identify segregation necessary for RISK CONTROL

The MANUFACTURER shall identify any segregation between SOFTWARE  ITEMS  that is necessary for RISK CONTROL, and state how to ensure that such segregation is effective. [Class   C]

NOTE   An  example of  segregation is  to have  SOFTWARE ITEMS  execute on  different processors. The effectiveness   of the segregation can be ensured by having no shared resources between the processors. Other means  of  segregation can be applied when effectiveness can be ensured by the software ARCHITECTURE design (see B.4.3).

    - Verify software ARCHITECTURE

The MANUFACTURER shall verify and document  that:

- the ARCHITECTURE of the software implements SYSTEM and software requirements including those relating to RISK CONTROL;

- the software ARCHITECTURE is able to support interfaces between SOFTWARE ITEMS and between SOFTWARE ITEMS and hardware;  and

- the MEDICAL DEVICE ARCHITECTURE supports proper operation of any SOUP items. [Class B, C]

NOTE       A TRACEABILITY analysis of ARCHITECTURE to software requirements can be used to satisfy requirement a).

#### * Software detailed design

    - Subdivide software into SOFTWARE UNITS

The MANUFACTURER shall subdvide the software until it is represented  by  SOFTWARE  UNITS. [Class B, C]

NOTE    Some SOFTWARE SYSTEMS  are not divided further.

    - Develop detailed design for each SOFTWARE  UNIT

The MANUFACTURER shall document a design with enough detail to allow correct implementation of each SOFTWARE UNIT. [Class  C]

#### Develop detailed design for interfaces

The MANUFACTURER  shall document a  design for any interfaces between the SOFTWARE  UNIT  and external components (hardware or software), as well as any interfaces between SOFTWARE UNITS, detailed enough to implement each SOFTWARE UNIT and its interfaces correctly. [Class   C]

#### Verify detailed design

– 24 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

The MANUFACTURER shall verify and document that the software detailed  design:

- implements the software ARCHITECTURE; and

- is free from contradiction with the software ARCHITECTURE. [Class C]

NOTE It is acceptable to use a TRACEABILITY analysis of ARCHITECTURE to software detailed design to satisfy requirement a).

  - * SOFTWARE UNIT  implementation

    - Implement each SOFTWARE UNIT

The MANUFACTURER shall implement each SOFTWARE UNIT. [Class A, B,   C]

    - Establish SOFTWARE UNIT VERIFICATION  PROCESS

The MANUFACTURER shall establish strategies, methods and procedures for verifying the SOFTWARE UNITS. Where VERIFICATION is done by testing, the test procedures  shall  be EVALUATED for adequacy. [Class B, C]

NOTE        It  is  acceptable to  combine integration testing and SOFTWARE SYSTEM  testing into a  single plan and set of

ACTIVITIES.

    - SOFTWARE UNIT acceptance criteria

The MANUFACTURER shall establish acceptance criteria for SOFTWARE UNITS prior to integration into larger SOFTWARE ITEMS as appropriate, and ensure that SOFTWARE UNITS meet acceptance criteria. [Class B, C]

NOTE     Examples of acceptance criteria  are:

- does the software code implement requirements including RISK CONTROL    measures?

- is the software code free from contradiction with the interface design of the SOFTWARE UNIT?

- does the software code conform to programming procedures or coding    standards?

    - Additional SOFTWARE UNIT acceptance  criteria

When present in the design, the MANUFACTURER shall include additional acceptance criteria as appropriate for:

- proper event sequence;

- data and control flow;

- planned resource allocation;

- fault handling (error definition, isolation, and recovery);

- initialisation of variables;

- self-diagnostics;

- memory management and memory overflows; and

- boundary conditions. [Class C]

    - SOFTWARE UNIT VERIFICATION

The MANUFACTURER shall perform the SOFTWARE UNIT VERIFICATION and document the results. [Class B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 25 –

+AMD1:2015 CSV  IEC 2015

#### * Software integration and integration testing

    - Integrate SOFTWARE UNITS

The MANUFACTURER shall integrate the SOFTWARE UNITS in accordance with the integration plan (see 5.1.5). [Class B, C]

#### Verify software integration

The MANUFACTURER shall verify that the SOFTWARE UNITS have been integrated into SOFTWARE ITEMS and/or the SOFTWARE SYSTEM  in  accordance with the integration plan (see 5.1.5) and retain records of the evidence of such  verification.

[Class B, C]

NOTE  This  VERIFICATION  is  only  that the integration has  been done according to the plan. This VERIFICATION  is   most likely implemented by some form of   inspection.

#### Software integration testing

The MANUFACTURER shall test the integrated SOFTWARE ITEMS in accordance with the integration plan (see 5.1.5) and document the results. [Class B,  C]

#### Software integration testing content

For  software  integration  testing,  the  MANUFACTURER   shall  address  whether  the    integrated

SOFTWARE ITEM performs as intended. [Class B, C]

NOTE 1  Examples to be considered  are:

- the required functionality of the  software;

- implementation of RISK CONTROL  measures;

- specified timing and other  behaviour;

- specified functioning of internal and external interfaces;  and

- testing under abnormal conditions including foreseeable   misuse.

NOTE 2      It  is  acceptable to  combine integration testing and SOFTWARE SYSTEM  testing into a  single plan and set of

ACTIVITIES.

#### EVALUATE  software integration test procedures

The MANUFACTURER shall EVALUATE the integration test procedures for adequacy. [Class B,   C]

#### Conduct regression tests

When software items are integrated, the MANUFACTURER shall conduct REGRESSION TESTING appropriate to demonstrate that defects have not been introduced into previously integrated software. [Class B, C]

#### Integration test record contents

The MANUFACTURER shall:

- document the test result (pass/fail and a list of  ANOMALIES);

- retain sufficient records to permit the test to be repeated;  and

- identify the tester. [Class B, C]

NOTE      Requirement b) could be implemented by retaining, for  example:

- test case specifications showing required actions and expected   results;

- records of the equipment;

– 26 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

records of the test environment (including software tools) used for    test.

#### Use software problem resolution PROCESS

The MANUFACTURER shall enter ANOMALIES found during software integration and integration testing into a software problem resolution PROCESS. [Class B,  C]

NOTE    See Clause 9.

  - * SOFTWARE SYSTEM  testing

#### Establish tests for software requirements

- The MANUFACTURER shall establish and perform a set of tests, expressed as input stimuli, expected outcomes, pass/fail criteria and procedures, for conducting SOFTWARE SYSTEM testing, such that all software requirements are covered. [Class A, B,  C]

NOTE 1 It  is  acceptable to  combine integration testing and  SOFTWARE SYSTEM  testing into a  single plan and  set of ACTIVITIES. It is also acceptable to test software requirements in earlier phases.

NOTE 2 Not only separate tests for each requirement, but also tests of combinations of requirements can be performed, especially if dependencies between requirements   exist.

- The MANUFACTURER shall EVALUATE the adequacy of VERIFICATION strategies and test procedures.

#### Use software problem resolution PROCESS

The MANUFACTURER shall enter ANOMALIES found during software system testing into a software problem resolution PROCESS. [Class A, B,  C]

#### Retest after changes

When changes are made during SOFTWARE SYSTEM testing, the MANUFACTURER   shall:

- repeat tests, perform modified tests or perform additional tests, as appropriate, to verify the effectiveness of the change in correcting the problem;

- conduct testing appropriate to demonstrate that unintended side effects have not been introduced; and

- perform relevant RISK MANAGEMENT ACTIVITIES as defined in 7.4. [Class A, B, C]

    - Evaluate SOFTWARE SYSTEM  testing

The MANUFACTURER shall EVALUATE the appropriateness of VERIFICATION strategies and test procedures.

The MANUFACTURER shall verify that:

- all software requirements have been tested or otherwise  VERIFIED;

- the TRACEABILITY between software requirements and tests or other VERIFICATION  is recorded; and

- test results meet the required pass/fail criteria. [Class A, B, C]

    - SOFTWARE SYSTEM test record contents

In order to support the repeatability of tests, the MANUFACTURER shall  document:

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 27 –

+AMD1:2015 CSV  IEC 2015

- a reference to test case procedures showing required actions and expected  results;

- the test result (pass/fail and a list of  ANOMALIES);

- the version of software tested;

- relevant hardware and software test configurations;

- relevant test tools;

- date tested; and

- the identity of the person responsible for executing the test and recording the test results. [Class A, B, C]

  - * Software RELEASE for utilization at a SYSTEM   level

    - Ensure software VERIFICATION is complete

The MANUFACTURER shall ensure that all software VERIFICATION ACTIVITIES have been completed and the results have been EVALUATED before the software is released. [Class A, B,   C]

    - Document known residual ANOMALIES

The MANUFACTURER shall document all known residual ANOMALIES. [Class A, B,   C]

    - EVALUATE  known residual ANOMALIES

The MANUFACTURER shall ensure that all known residual ANOMALIES have been EVALUATED to ensure that they do not contribute to an unacceptable RISK. [Class B,   C]

    - Document released VERSIONS

The MANUFACTURER shall document the VERSION of the MEDICAL DEVICE SOFTWARE that is being released. [Class A, B, C]

#### Document how released software was created

The MANUFACTURER shall document the procedure and environment used to create the released software. [Class B, C]

#### Ensure activities and tasks are complete

The  MANUFACTURER  shall  ensure  that  all  software  development  plan  (or  maintenance plan)

ACTIVITIES and TASKS are complete along with the associated documentation. [Class B,   C]

NOTE    See 5.1.3.b).

#### Archive software

The MANUFACTURER shall archive:

- the MEDICAL DEVICE  SOFTWARE and CONFIGURATION ITEMS;  and

- the documentation

for at least a period of time determined as the longer of: the life time of the MEDICAL DEVICE SOFTWARE as defined by the MANUFACTURER or a time specified by relevant regulatory requirements. [Class A, B, C]

#### Assure reliable delivery of released software

The MANUFACTURER shall establish procedures to ensure that the released MEDICAL DEVICE SOFTWARE  can  be  reliably  delivered  to  the  point  of  use  without  corruption  or  unauthorised

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 28 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

change. These procedures shall address the production and handling of media containing the

MEDICAL DEVICE  SOFTWARE including as appropriate:

- replication,

- media labelling,

- packaging,

- protection,

- storage, and

- delivery. [Class A, B, C]

- Software maintenance PROCESS

#### * Establish software maintenance plan

The  MANUFACTURER  shall  establish a  software maintenance plan  (or  plans)  for conducting the

ACTIVITIES and TASKS of the maintenance PROCESS.  The plan shall address the   following:

- procedures for:

  - receiving,

  - documenting,

  - evaluating,

  - resolving and

  - tracking

feedback arising after release of the MEDICAL DEVICE   SOFTWARE;

- criteria for determining whether feedback is considered to be a  problem;

- use of the software RISK MANAGEMENT  PROCESS;

- use of the software problem resolution PROCESS for analysing  and  resolving  problems arising after release of the MEDICAL DEVICE  SOFTWARE;

- use of the software configuration management PROCESS (Clause 8) for managing modifications to the existing SOFTWARE SYSTEM;  and

- procedures to EVALUATE and  implement:

  - upgrades,

  - bug fixes,

  - patches and

  - obsolescence of SOUP.

[Class A, B, C]

#### * Problem and modification analysis

    - Document and EVALUATE  feedback

#### Monitor feedback

The MANUFACTURER shall monitor feedback on MEDICAL DEVICE SOFTWARE released for intended use.

[Class A, B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 29 –

+AMD1:2015 CSV  IEC 2015

      - Document and EVALUATE  feedback

Feedback shall be documented and EVALUATED to determine whether a problem exists in a released MEDICAL DEVICE SOFTWARE. Any such problem shall be recorded as a PROBLEM REPORT (see Clause 9). PROBLEM REPORTS shall include actual or potential adverse events, and deviations from specifications. [Class A, B,  C]

      - Evaluate PROBLEM REPORT’S affects on  SAFETY

Each PROBLEM REPORT shall be EVALUATED to determine how it affects the SAFETY of MEDICAL DEVICE SOFTWARE released for intended use (see 9.2) and whether a change to that software is needed to address the problem. [Class A, B,  C]

#### Use software problem resolution PROCESS

The MANUFACTURER shall use the software problem resolution PROCESS (see Clause  9)  to address PROBLEM REPORTS. [Class A, B,  C]

NOTE A problem could show that a SOFTWARE SYSTEM or SOFTWARE  ITEM  has  not  been  placed  in  the  correct software safety class. The problem resolution process can suggest changes of the software safety class. When the PROCESS has been completed, any change of safety class in the SOFTWARE SYSTEM or its SOFTWARE ITEMS should be made known and documented.

    - Analyse CHANGE REQUESTS

In addition to the analysis required by Clause 9, the MANUFACTURER shall analyse each CHANGE REQUEST for its effect on the organization, MEDICAL DEVICE SOFTWARE released for intended use , and SYSTEMS with which it interfaces. [Class A, B,  C]

    - CHANGE REQUEST approval

The  MANUFACTURER  shall  EVALUATE   and  approve  CHANGE  REQUESTS   which  modify  released

MEDICAL DEVICE  SOFTWARE. [Class A, B, C]

    - Communicate to users and regulators

The MANUFACTURER shall identify the approved CHANGE REQUESTS that affect released MEDICAL DEVICE SOFTWARE.

As required by local regulation, the MANUFACTURER shall inform users and regulators   about:

- any problem in released MEDICAL DEVICE SOFTWARE and the consequences of continued unchanged use; and

- the nature of any available changes to released MEDICAL DEVICE  SOFTWARE  and how to  obtain and install the changes.

[Class A, B, C]

#### * Modification implementation

    - Use established PROCESS to implement modification

The MANUFACTURER shall identify and perform any Clause 5 ACTIVITIES that need to be repeated as a result of the modification. [Class A, B,  C]

NOTE     For requirements relating to RISK MANAGEMENT of software changes    see 7.4.

    - Re-release modified SOFTWARE SYSTEM

The MANUFACTURER shall release modifications according to 5.8. [Class A, B,   C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 30 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

NOTE Modifications can be released as part of a full re-release of a SOFTWARE SYSTEM or as a modification kit comprising changed SOFTWARE ITEMS and the necessary tools to install the changes as modifications to an existing SOFTWARE SYSTEM.

- * Software RISK MANAGEMENT  PROCESS

#### * Analysis of software contributing to hazardous  situations

    - Identify SOFTWARE ITEMS that could contribute to a hazardous  situation

The MANUFACTURER shall identify SOFTWARE ITEMS that could contribute to a hazardous situation identified in the MEDICAL DEVICE RISK ANALYSIS ACTIVITY of ISO 14971 (see 4.2). [Class B, C]

NOTE The hazardous situation could be the direct result of software failure or the result of the failure of a  RISK CONTROL  measure that is implemented in  software.

#### Identify potential causes of contribution to a hazardous  situation

The MANUFACTURER shall identify potential causes of the SOFTWARE ITEM identified above contributing to a hazardous situation.

The MANUFACTURER shall consider potential causes including, as  appropriate:

- incorrect or incomplete specification of functionality;

- software defects in the identified SOFTWARE ITEM  functionality;

- failure or unexpected results from  SOUP;

- hardware failures or other software defects that could result in unpredictable software operation; and

- reasonably foreseeable misuse. [Class B, C]

    - EVALUATE  published SOUP  ANOMALY lists

If failure or unexpected results from SOUP is a potential cause  of the  SOFTWARE  ITEM  contributing to a hazardous situation, the MANUFACTURER shall EVALUATE as a minimum any ANOMALY  list published by the supplier of the SOUP  item  relevant to the VERSION  of the SOUP  item used in the MEDICAL DEVICE to determine if any of the known ANOMALIES  result  in  a  sequence of events that could result in a hazardous situation. [Class B,   C]

#### Document potential causes

The  MANUFACTURER   shall  document  in  the  RISK  MANAGEMENT  FILE   potential  causes  of    the

SOFTWARE ITEM contributing to a hazardous situation (see ISO 14971). [Class B,   C]

  - RISK CONTROL measures

    - Define RISK CONTROL measures

For each case documented in the RISK MANAGEMENT FILE where a SOFTWARE ITEM  could  contribute to a HAZARDOUS SITUATION, the MANUFACTURER shall define and document RISK CONTROL measures in accordance with ISO 14971. [Class B,  C]

NOTE The RISK CONTROL measures can be implemented in hardware, software, the working environment or user instruction.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 31 –

+AMD1:2015 CSV  IEC 2015

    - RISK CONTROL measures implemented in  software

If a RISK CONTROL measure is implemented as part of the functions of a SOFTWARE ITEM, the

MANUFACTURER shall:

- include the RISK CONTROL measure in the software  requirements;

- assign to each SOFTWARE ITEM that contributes to the implementation of a RISK CONTROL measure a software safety class based on the RISK that the RISK CONTROL measure is controlling (see 4.3 a)); and

- develop the SOFTWARE ITEM in accordance with Clause 5. [Class B, C]

NOTE        This requirement provides additional detail for RISK CONTROL requirements of ISO 14971

  - VERIFICATION of RISK CONTROL  measures

    - Verify RISK CONTROL measures

The implementation of each RISK CONTROL measure documented in 7.2 shall be VERIFIED, and this VERIFICATION shall be documented. The MANUFACTURER shall review the RISK CONTROL measure and determine if it could result in a new HAZARDOUS SITUATION. [Class B,   C]

#### 7.3.2

Not used.

7.3.3   Document TRACEABILITY

The MANUFACTURER shall document TRACEABILITY of software HAZARDS as   appropriate:

- from the hazardous situation to the SOFTWARE  ITEM;

- from the SOFTWARE ITEM to the specific software  cause;

- from the software cause to the RISK CONTROL measure;  and

- from the RISK CONTROL measure to the VERIFICATION of the RISK CONTROL measure. [Class B, C]

NOTE    See ISO 14971 – RISK MANAGEMENT  report.

  - RISK MANAGEMENT of software  changes

    - Analyse changes to MEDICAL DEVICE SOFTWARE with respect to  SAFETY

The MANUFACTURER shall analyse changes to the MEDICAL DEVICE SOFTWARE (including SOUP) to determine whether:

- additional potential causes are introduced contributing to a hazardous situation;  and

- additional software RISK CONTROL measures are required. [Class A, B, C]

    - Analyse impact of software changes on existing RISK CONTROL  measures

The MANUFACTURER shall analyse changes to the software, including changes to SOUP, to determine whether the software modification could interfere with existing RISK CONTROL measures. [Class B, C]

    - Perform RISK MANAGEMENT ACTIVITIES based on  analyses

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 32 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

The MANUFACTURER  shall perform relevant RISK MANAGEMENT  ACTIVITIES  defined in 7.1, 7.2   and

7.3 based on these analyses. [Class B,  C]

### * Software configuration management  PROCESS

#### * Configuration identification

    - Establish means to identify CONFIGURATION  ITEMS

The MANUFACTURER shall establish a scheme for the unique identification of CONFIGURATION ITEMS and their VERSIONS to be controlled according to the development and configuration planning specified in 5.1. [Class A, B,  C]

    - Identify SOUP

For each SOUP  CONFIGURATION ITEM being used, including standard libraries, the  MANUFACTURER

shall document:

- the title,

- the MANUFACTURER, and

- the unique SOUP designator [Class A, B, C]

NOTE The unique SOUP designator could be, for example, a  VERSION,  a  release  date,  a  patch  number  or  an upgrade designation.

#### Identify SYSTEM configuration documentation

The MANUFACTURER shall document the set of CONFIGURATION ITEMS and their VERSIONS that comprise the SOFTWARE SYSTEM configuration. [Class A, B,  C]

#### 8.2   * Change control

8.2.1   Approve CHANGE REQUESTS

The MANUFACTURER  shall change CONFIGURATION  ITEMS  identified to be controlled according   to

8.1 only in response to an approved CHANGE REQUEST. [Class A, B,   C]

NOTE 1 The decision to approve a CHANGE REQUEST can be integral to  the  change  control  PROCESS  or  part  of another PROCESS. This subclause only requires that approval of a change precede its implementation.

NOTE 2 Different acceptance PROCESSES can be used for CHANGE REQUESTS at different stages of the life cycle, as stated in plans, see 5.1.1  d) and 6.1   e).

#### Implement changes

The MANUFACTURER shall implement the change as specified in the CHANGE REQUEST. The MANUFACTURER shall identify and perform any ACTIVITY that needs to be repeated as a result of the change, including changes to the software safety classification of SOFTWARE SYSTEMS and SOFTWARE ITEMS. [Class A, B,  C]

NOTE   This subclause states how the change should be implemented to achieve adequate change control. It does    not imply that the implementation is an integral part of the  change control PROCESS. Implementation should use  planned PROCESSES, see 5.1.1 e) and 6.1  e).

#### Verify changes

The MANUFACTURER shall verify the change, including repeating any VERIFICATION that has been invalidated by the change and taking into account 5.7.3 and 9.7. [Class A, B,   C]

NOTE	This  subclause only requires that changes be  VERIFIED. It  does  not  imply that VERIFICATION  is    an integral part of the change control PROCESS. VERIFICATION should use planned PROCESSES, see 5.1.1 e) and 6.1 e).

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 33 –

+AMD1:2015 CSV  IEC 2015

    - Provide means for TRACEABILITY of  change

The MANUFACTURER shall maintain records of the relationships and dependencies  between:

- CHANGE REQUEST;

- relevant PROBLEM REPORT; and

- approval of the CHANGE  REQUEST

[Class A, B, C]

#### 8.3   * Configuration status accounting

The MANUFACTURER shall retain retrievable records of the history of controlled CONFIGURATION ITEMS including SYSTEM configuration. [Class A, B,  C]

- * Software problem resolution  PROCESS

  - Prepare PROBLEM REPORTS

The MANUFACTURER shall prepare a PROBLEM REPORT for each problem detected in the MEDICAL DEVICE SOFTWARE. PROBLEM REPORTS shall include a statement of criticality (for example, effect on performance, SAFETY, or SECURITY) as well as other information that may aid  in  the resolution of the problem (for example, devices affected, supported accessories  affected).

[Class A, B, C]

NOTE      Problems can be discovered before or after release, inside the MANUFACTURER’S  organization or outside   it.

#### Investigate the problem

The MANUFACTURER shall:

- investigate the problem and if possible identify the  causes;

- EVALUATE  the problem’s relevance to SAFETY  using the software     RISK  MANAGEMENT  PROCESS

(Clause 7);

- document the outcome of the investigation and evaluation;  and

- create a CHANGE REQUEST(S) for actions needed to correct the problem, or document the rationale for taking no action.

[Class A, B, C]

NOTE A problem does not have to be corrected for the MANUFACTURER to  comply  with  the  software  problem resolution PROCESS, provided that the problem is not relevant to   SAFETY.

#### Advise relevant parties

The MANUFACTURER shall advise relevant parties of the existence of the problem,  as  appropriate.

[Class A, B, C]

NOTE   Problems can be discovered before or after release, inside the MANUFACTURER’S  organisation or outside it.    The MANUFACTURER  determines the relevant parties depending on the   situation.

#### Use change control process

The MANUFACTURER shall approve and implement all CHANGE REQUESTS, observing the requirements of the change control PROCESS (see 8.2). [Class A, B,   C]

#### Maintain records

– 34 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

The MANUFACTURER shall maintain records of PROBLEM REPORTS and their resolution including their VERIFICATION.

The MANUFACTURER shall update the RISK MANAGEMENT FILE as appropriate. [Class A, B,    C]

#### Analyse problems for trends

The MANUFACTURER shall perform    analysis to detect trends in PROBLEM REPORTS. [Class A, B, C]

#### Verify software problem resolution

The MANUFACTURER shall verify resolutions to determine  whether:

- problem has been resolved and the PROBLEM REPORT has been  closed;

- adverse trends have been reversed;

- CHANGE REQUESTS  have been implemented in the appropriate MEDICAL DEVICE  SOFTWARE   and

ACTIVITIES; and

- additional problems have been introduced. [Class A, B, C]

#### Test documentation contents

When testing, retesting or REGRESSION TESTING SOFTWARE ITEMS and SYSTEMS  following  a change, the MANUFACTURER shall include in the test  documentation:

- test results;

- ANOMALIES found;

- the VERSION of software tested;

- relevant hardware and software test configurations;

- relevant test tools;

- date tested; and

- identification of the tester. [Class A, B, C]

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 35 –

+AMD1:2015 CSV  IEC 2015

# Annex A

## (informative)

Rationale for the requirements of this  standard

Rationale for the clauses of this standard is provided in this  annex.

### Rationale

The primary requirement of this standard is that a set of PROCESSES be followed in the development and maintenance of MEDICAL DEVICE SOFTWARE, and that the choice of PROCESSES be appropriate to the RISKS to the patient and other people. This follows from the belief that testing of software is not sufficient to determine that it is safe in   operation.

The PROCESSES required by this standard fall into two  categories:

- PROCESSES which are required to determine the RISKS  arising from the operation of each

SOFTWARE ITEM in the software;

- PROCESSES which are required to achieve an appropriately low probability of software failure for each SOFTWARE ITEM, chosen on the basis of these determined   RISKS.

This standard requires the first category to be performed for all MEDICAL DEVICE SOFTWARE  and   the second category to be performed for selected SOFTWARE  ITEMS.

A claim of compliance with this standard should therefore include a documented RISK ANALYSIS that identifies foreseeable sequences of events that include software and that can result in a HAZARDOUS SITUATION (see ISO 14971). HAZARDOUS SITUATIONS that can be indirectly caused by software (for example, by providing misleading information that could cause inappropriate treatment to be administered) should be included in this RISK   ANALYSIS.

All ACTIVITIES that are required as part of the first category of PROCESSES are identified in the normative text as "[Class A, B, C]", indicating that they are required irrespective of the classification of the software to which they  apply.

ACTIVITIES are required for all classes A, B, and C for the following   reasons:

- the ACTIVITY produces a plan relevant to RISK MANAGEMENT or software   safety classification;

- the ACTIVITY produces an output that is an input to RISK MANAGEMENT or software safety classification;

- the ACTIVITY is a part of RISK MANAGEMENT or software safety   classification;

- the ACTIVITY establishes an administration  system,  documentation  or record-keeping system that supports RISK MANAGEMENT or software safety  classification;

- the ACTIVITY normally takes place when the classification of the related  software  is unknown;

- the ACTIVITY can cause a change that could invalidate the current software safety classification of the associated software. This includes the discovery and analysis of safety related problems after release.

Other PROCESSES are required only for SOFTWARE SYSTEMS or SOFTWARE ITEMS classified in software safety classes B or C. ACTIVITIES  required as parts of these PROCESSES  are identified  in the normative text as "[Class B, C]", or "[Class C]" indicating that they  are  required selectively depending on the classification of the software to which they  apply.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 36 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

ACTIVITIES are required selectively for software in classes B and C for the following   reasons:

- the ACTIVITY enhances the reliability of the software by requiring more detail or more rigor in the design, testing or other VERIFICATION;

- the ACTIVITY is an administrative ACTIVITY that supports another  ACTIVITY  required  for  classes B or C;

- the ACTIVITY supports the correction of safety-related  problems;

- the ACTIVITY produces records of design, implementation, VERIFICATION and release of safety-related software.

ACTIVITIES are required selectively for software in class C for the following   reasons:

- the ACTIVITY further enhances the reliability of  the  software  by requiring more detail, or more rigour, or attention to specific issues in the design, testing or other   VERIFICATION

Note that all PROCESSES and ACTIVITIES defined in this standard are considered valuable in assuring the development and maintenance of high quality software. The omission of many of these PROCESSES and ACTIVITIES as requirements for software in class A should not imply that these PROCESSES and ACTIVITIES would not be of value or  are  not  recommended.  Their  omission is intended to recognize that software that cannot cause a HAZARD can be assured of SAFETY and effectiveness primarily through overall validation ACTIVITY during the design of a MEDICAL DEVICE (which is outside the scope of this standard) and through some simple software life cycle controls.

### Summary of requirements by class

Table A.1 summarizes which software safety classes are assigned to each requirement. This table is informative and only provided for convenience. The normative section identifies the software safety classes for each requirement.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 37 –

+AMD1:2015 CSV  IEC 2015

#### Table A.1 – Summary of requirements by software safety  class

| Clauses and subclauses |  |  |  |  | Class A | Class B | Class C |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Clause | 4 | All requirements |  |  | X | X | X |
| 5.1 |  | 5.1.1, 5.1.2, 5.1.3, 5.1.6, 5.1.7, | 5.1.8, | 5.1.9 | X | X | X |
| 5.1.5, 5.1.10, 5.1.11, 5.1.12 |  |  |  |  |  | X | X |
| 5.1.4 |  |  |  |  |  |  | X |
| 5.2 |  | 5.2.1, 5.2.2, 5.2.4, 5.2.5, 5.2.6 |  |  | X | X | X |
| 5.2.3 |  |  |  |  |  | X | X |
| 5.3 |  | 5.3.1, 5.3.2, 5.3.3, 5.3.4, 5.3.6 |  |  |  | X | X |
| 5.3.5 |  |  |  |  |  |  | X |
| 5.4 |  | 5.4.1 |  |  |  | X | X |
| 5.4.2, 5.4.3, 5.4.4 |  |  |  |  |  |  | X |
| 5.5 |  | 5.5.1 |  |  | X | X | X |
| 5.5.2, 5.5.3, 5.5.5 |  |  |  |  |  | X | X |
| 5.5.4 |  |  |  |  |  |  | X |
| 5.6 |  | All requirements |  |  |  | X | X |
| 5.7 |  | All requirements |  |  | X | X | X |
| 5.8 |  | 5.8.1, 5.8.2, 5.8.4, 5.8.7, 5.8.8 |  |  | X | X | X |
| 5.8.3, 5.8.5, 5.8.6 |  |  |  |  |  | X | X |
| Clause | 6 | All requirements |  |  | X | X | X |
| 7.1 |  | All requirements |  |  |  | X | X |
| 7.2 |  | All requirements |  |  |  | X | X |
| 7.3 |  | All requirements |  |  |  | X | X |
| 7.4 |  | 7.4.1 |  |  | X | X | X |
| 7.4.2, 7.4.3 |  |  |  |  |  | X | X |
| Clause | 8 | All requirements |  |  | X | X | X |
| Clause | 9 | All requirements |  |  | X | X | X |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 38 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

# Annex B

## (informative)

Guidance on the provisions of this  standard

### Scope

#### Purpose

The purpose of this standard is to provide a development PROCESS that will  consistently  produce high quality, safe MEDICAL DEVICE SOFTWARE. To accomplish  this,  the  standard identifies the minimum ACTIVITIES and TASKS that need to be accomplished to provide confidence that the software has been developed in a manner that is likely to produce highly reliable and safe MEDICAL DEVICE  SOFTWARE.

This annex provides guidance for the application of the requirements of this standard.  It does not add to, or otherwise change, the requirements of this standard. This annex can be used to better understand the requirements of this  standard.

Note that in this standard, ACTIVITIES are subclauses called  out  within  the  PROCESSES  and TASKS are defined within the ACTIVITIES. For example, the ACTIVITIES defined for the software development PROCESS are software development planning, software requirements analysis, software ARCHITECTURAL design, software detailed design, SOFTWARE UNIT implementation and VERIFICATION, software integration and integration testing, SOFTWARE SYSTEM testing, and software release. The TASKS within these ACTIVITIES are the individual   requirements.

This standard does not require a particular SOFTWARE DEVELOPMENT  LIFE  CYCLE  MODEL.  However, compliance with this standard does imply dependencies between PROCESSES,  because inputs of a PROCESS are generated by another PROCESS. For example, the software safety classification of the SOFTWARE SYSTEM should be completed after the RISK ANALYSIS PROCESS has established what HARM could arise from failure of the   SOFTWARE SYSTEM.

Because of such logical dependencies between processes, it is easiest to describe the processes in this standard in a sequence, implying a “waterfall” or “once-through” life cycle model. However, other life cycles can also be used. Some development (model) strategies as defined at ISO/IEC 12207 [9] include (see also Table  B.1):

- Waterfall. The “once-through" strategy, also called “waterfall”, consists of performing the development PROCESS a single time. Simplistically: determine customer needs, define requirements, design the SYSTEM, implement the system, test, fix and  deliver.

- Incremental: The “incremental” strategy determines customer needs and defines the  SYSTEM requirements, then performs the rest of the development in a sequence of builds.  The first build incorporates part of the planned capabilities, the next build adds more capabilities, and so on, until the SYSTEM is  complete.

- Evolutionary: The “evolutionary” strategy also develops a SYSTEM in builds but differs from  the incremental strategy in acknowledging that the user need is not fully understood and all requirements cannot be defined up front. In this strategy, customer needs and SYSTEM requirements are partially defined up front, then are refined in each succeeding  build.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 39 –

+AMD1:2015 CSV  IEC 2015

#### Table B.1 – Development (model) strategies as defined in ISO/IEC  12207

| Development Strategy | Define all requirements first? | Multiple development cycles? | Distribute interim software? |
| --- | --- | --- | --- |
| Waterfall (Once-through) | yes | no | no |
| Incremental (Preplanned product improvement) | yes | yes | maybe |
| Evolutionary | no | yes | yes |

Whichever life cycle is chosen it is necessary to maintain the logical dependencies between PROCESS outputs such as  specifications, design documents and software. The waterfall life  cycle model achieves this by delaying the start of a  PROCESS  until the inputs for that PROCESS are complete and approved.

Other life cycles, particularly evolutionary life cycles, permit PROCESS outputs to be produced before all the inputs for that PROCESS are available. For example, a new SOFTWARE ITEM can be specified, classified, implemented and VERIFIED before the whole software ARCHITECTURE has been finalised. Such life cycles carry the RISK that a change or development in one PROCESS output will invalidate another PROCESS output. All life cycles therefore use a comprehensive configuration management system to ensure that all PROCESS outputs are brought to  a  consistent state and the dependencies maintained.

The following principles are important regardless of the software development life cycle   used:

- All PROCESS outputs should be maintained in a consistent state; whenever any PROCESS output is created or changed, all related PROCESS outputs should be updated promptly to maintain their consistency with each other and to maintain all dependencies explicitly or implicitly required by this standard;

- all PROCESS outputs should be available when needed as input to further work on  the software.

- before any MEDICAL DEVICE SOFTWARE is released, all PROCESS outputs should be consistent with each other and all dependencies between PROCESS outputs explicitly or implicitly required by this standard should be observed.

#### Field of application

This standard applies to the development and maintenance of  MEDICAL  DEVICE  SOFTWARE  as well as the development and maintenance of a MEDICAL DEVICE that includes   SOUP.

The use of this standard requires the MANUFACTURER to perform MEDICAL DEVICE RISK  MANAGEMENT that is compliant with ISO 14971. Therefore, when the MEDICAL DEVICE SYSTEM ARCHITECTURE includes an acquired component (this could be a purchased component or a component of unknown provenance), such as a printer/plotter that includes SOUP, the acquired component becomes the responsibility of the MANUFACTURER and must be included in the RISK MANAGEMENT of the MEDICAL DEVICE. It is assumed that through proper performance of MEDICAL DEVICE RISK MANAGEMENT, the MANUFACTURER would understand the component and recognize that it includes SOUP. The MANUFACTURER using this standard would invoke the software RISK MANAGEMENT PROCESS as part of the overall   MEDICAL DEVICE RISK MANAGEMENT PROCESS.

The maintenance of released MEDICAL DEVICE SOFTWARE applies to the post-production experience with the MEDICAL DEVICE SOFTWARE. Software maintenance includes the combination of all technical and administrative means, including supervision actions, to act on problem reports to retain an item in, or restore it to, a state in which it can perform a required function     as well as modification requests related to released MEDICAL  DEVICE  SOFTWARE. For example, this includes problem rectification, regulatory reporting, re-validation and preventive action. See ISO/IEC 14764 [10].

### Normative references

– 40 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

ISO/IEC 90003 [15] provides guidance for applying a quality management system to software development. This guidance is not required by this standard but is highly  recommended.

### Terms and definitions

Where possible, terms have been defined using definitions from international  standards.

This standard chose to use three terms to describe the decomposition of a SOFTWARE SYSTEM (top level). The SOFTWARE SYSTEM can be a subsystem of the MEDICAL DEVICE (see IEC 60601-  1-4 [2]) or a MEDICAL DEVICE in its  own  right, which then becomes a  software MEDICAL  DEVICE. The lowest level that is not further decomposed for the purposes of testing or software configuration management is the SOFTWARE UNIT. All levels of  composition, including the top and bottom levels, can be called SOFTWARE ITEMS. A SOFTWARE SYSTEM, then, is composed of one or more SOFTWARE ITEMS, and each SOFTWARE ITEM is composed of one or more SOFTWARE UNITS or decomposable SOFTWARE ITEMS. The responsibility is left to the  MANUFACTURER  to provide the definition and granularity of the SOFTWARE ITEMS and  SOFTWARE  UNITS.  Leaving these terms vague allows one to apply them to the many different development methods and types of software used in MEDICAL  DEVICES.

### General requirements

There is no known method to guarantee 100 % SAFETY for any kind of   software.

There are three major principles which promote SAFETY for MEDICAL DEVICE    SOFTWARE:

- RISK MANAGEMENT;

- quality management;

- software engineering.

For the development and maintenance of safe MEDICAL DEVICE SOFTWARE it is necessary to establish RISK MANAGEMENT as an integral part of a quality management system as an overall framework for the application of  appropriate  software  engineering methods and  techniques. The combination of these three concepts allows a MEDICAL DEVICE MANUFACTURER to follow a clearly structured and consistently repeatable decision-making PROCESS to promote SAFETY for MEDICAL DEVICE SOFTWARE.

#### Quality management system

A disciplined and effective set of software PROCESSES includes organizational PROCESSES  such  as  management, infrastructure, improvement, and training. To avoid duplication and to focus  this standard on software engineering, these PROCESSES have been omitted from this standard. These PROCESSES are covered by a quality management system. ISO 13485 [8] is an International Standard that is specifically intended for applying the concepts of quality management to MEDICAL DEVICES. Conformance to ISO 13485 quality management system requirements does not automatically constitute conformity with national or regional regulatory requirements. It is the MANUFACTURER’S responsibility to identify and establish compliance with relevant regulatory requirements.

    - RISK MANAGEMENT

Software development participates in RISK MANAGEMENT ACTIVITIES sufficiently to ensure that all reasonably foreseeable RISKS associated with the MEDICAL DEVICE SOFTWARE are   considered.

Rather than trying to define an appropriate RISK MANAGEMENT PROCESS in this software engineering standard, it is required that the MANUFACTURER  apply a RISK MANAGEMENT     PROCESS

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 41 –

+AMD1:2015 CSV  IEC 2015

that is compliant with ISO 14971, which deals explicitly with RISK MANAGEMENT for MEDICAL DEVICES. Specific software RISK MANAGEMENT ACTIVITIES  resulting from HAZARDOUS  SITUATIONS that have software as a contributing cause are identified in a supporting PROCESS described in Clause 7.

#### Software safety classification

The RISK associated with software as a part of a MEDICAL DEVICE, as an accessory to a MEDICAL DEVICE, or as a MEDICAL DEVICE in its own right, is used as the input to a software safety classification scheme, which then determines the PROCESSES to be used during  the development and maintenance of software.

RISK is considered to be a combination of the severity of HARM and the probability of its occurrence. However, no consensus exists for a method of quantitatively estimating the probability of occurrence of a software failure. When software is present in a sequence or combination of events leading to a HAZARDOUS SITUATION, the probability of the software failure occurring cannot be considered in estimating the RISK for the HAZARDOUS SITUATION. In such cases, considering a worst case probability is appropriate, and the probability for the software failure occurring should be set to 1. When it is possible to estimate the probability for the remaining events in the sequence (as it may be if they are not software) that probability can be used for the probability of the HAZARDOUS SITUATION occurring (P1  in Figure  B.2).

In many cases however, it might not be possible to estimate the probability for the remaining events in the sequence, and the RISK should be EVALUATED  on the basis of the nature of the HARM alone (the probability of the HAZARDOUS SITUATION occurring should be set to 1). RISK ESTIMATION in these cases should be focused on the SEVERITY of the HARM resulting from the HAZARDOUS SITUATION. Subjective rankings of probability can also be assigned based on clinical knowledge to distinguish failures that a clinician would be likely to detect from those that would not be detected and would be more likely to cause  HARM.

Estimates of probability of a HAZARDOUS SITUATION leading to HARM (P2 in Figure B.2) generally require clinical knowledge to distinguish between HAZARDOUS SITUATIONS where clinical practice would be likely to prevent HARM, and HAZARDOUS SITUATIONS that would be more likely to cause HARM.

– 42 –	IEC 62304:2006

Sequence of events

+AMD1:2015 CSV  IEC 2015

NOTE	P1 is the probability of a hazardous situation   occurring

P2 is the probability of a hazardous situation leading to    harm

Hazard

Exposure (P1)

Hazardous

situation

P2

Harm

Severity of

the harm

Probability

of occurrence of harm

Risk

P1  × P2

IEC

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

Figure B.2 – Pictorial representation of the relationship of HAZARD, sequence of  events,

HAZARDOUS SITUATION, and HARM  – from ISO 14971:2007 Annex  E

If a SOFTWARE SYSTEM is decomposed into SOFTWARE ITEMS, then  each SOFTWARE  ITEM  can  have its own software safety classification.

It is only possible to determine the RISK associated with failure of a   SOFTWARE ITEM:

- if a SYSTEM ARCHITECTURE and a software ARCHITECTURE  define the role of  the SOFTWARE  ITEM in terms of its purpose and its interfaces with other software and hardware   items;

- if changes to the SYSTEM are  controlled;

- after RISK ANALYSIS has been done on the ARCHITECTURE and RISK CONTROL measures specified.

This standard requires the minimum number of ACTIVITIES that will  achieve  the  above  conditions for all classes of software.

The end of the software ARCHITECTURE ACTIVITY  is  the earliest point in the development when  the full set of SOFTWARE ITEMS is defined and the RISK MANAGEMENT ACTIVITY has identified how the SOFTWARE ITEMS relate to SAFETY. This is therefore the earliest point at  which  SOFTWARE ITEMS can be classified definitively according to their SAFETY   role.

This point corresponds to the point where RISK CONTROL is begun in ISO   14971.

Before this point, the RISK MANAGEMENT PROCESS identifies ARCHITECTURAL RISK CONTROL measures, for example adding protective subsystems, or reducing the opportunities  for  software failures to cause HARM. After this point, the RISK MANAGEMENT  PROCESS  uses PROCESSES aimed at reducing the probability of failure of SOFTWARE ITEMS. In other words, the classification of a SOFTWARE ITEM specifies PROCESS-based RISK CONTROL measures to be  applied to that item.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 43 –

+AMD1:2015 CSV  IEC 2015

It is expected that MANUFACTURERS will find it useful to classify software before this point, for example to focus attention on areas to be investigated, but such classification should be regarded as preliminary and should not be used to justify the omission of  PROCESSES.

The software safety classification scheme is not intended to align with the RISK  classifications    of ISO 14971. Whereas the ISO 14971 scheme classifies RISK according to their severity and likelihood, the software safety classification scheme classifies SOFTWARE SYSTEMS and SOFTWARE ITEMS according to the PROCESSES to be applied in their development and maintenance.

As the design evolves, new RISKS might become evident. Therefore, RISK  MANAGEMENT  should  be applied as an integral part of the development PROCESS. This permits the development of an ARCHITECTURAL design that identifies a complete set  of  SOFTWARE  ITEMS, including those that are required to function correctly to assure safe operation and those that prevent faults from causing HARM.

The software ARCHITECTURE should promote segregation of software items that are required for safe operation and should describe the methods used to ensure effective segregation of those SOFTWARE ITEMS. Segregation is not restricted to physical (processor or memory partition) separation but includes any mechanism that prevents one SOFTWARE ITEM from negatively affecting another. The adequacy of a segregation is determined based on the RISKS  involved  and the rationale which is required to be  documented.

As stated in B.3, this standard chooses to use three terms to describe the decomposition of a

SOFTWARE SYSTEM (top level).

Figure B.1 illustrates the possible partitioning for SOFTWARE ITEMS  within a  SOFTWARE  SYSTEM and how the software safety classes would be applied to the group of SOFTWARE ITEMS in the decomposition.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 44 –	IEC 62304:2006

SOFTWARE SYSTEM /

SOFTWARE ITEM

(CLASS C)

SOFTWARE ITEM

X

(Class A)

SOFTWARE ITEM

Y

(Class C)

SOFTWARE ITEM

W

(Class B)

SOFTWARE ITEM

Z

(Class C)

IEC  724/06

+AMD1:2015 CSV  IEC 2015

Figure B.1 – Example of partitioning of SOFTWARE  ITEMS

For this example, the MANUFACTURER knows, due to  the  type  of  MEDICAL  DEVICE  SOFTWARE being developed, that the preliminary software safety classification for the SOFTWARE SYSTEM is software safety class C. During software ARCHITECTURE design the MANUFACTURER has decided  to partition the SYSTEM, as shown, with 3 SOFTWARE ITEMS – X, W and Z. The MANUFACTURER is able to segregate all SOFTWARE SYSTEM contributions to HAZARDOUS  SITUATIONS  which  could result in death or SERIOUS INJURY to SOFTWARE ITEM Z and all remaining SOFTWARE SYSTEM contributions to HAZARDOUS SITUATIONS which could result in a  non-SERIOUS  INJURY  to  SOFTWARE ITEM W. SOFTWARE ITEm W is classified as software safety class B  and SOFTWARE  ITEM Z is at software safety class C. SOFTWARE ITEM Y therefore must be classified as Class C, per 4.3 d). The SOFTWARE SYSTEM is also at a software safety class C per this requirement. SOFTWARE ITEM X has been classified at a  software safety class of  A.  The MANUFACTURER  is able to document a rationale for the segregation between SOFTWARE ITEMS X and Y, as well as SOFTWARE ITEMS W and Z, to assure the integrity of the segregation. If  segregation  is  not possible between SOFTWARE ITEMS X and Y, then SOFTWARE ITEM X must be  classified  in  software safety class C.

    - LEGACY SOFTWARE

Subclause 4.4 establishes a process for application of this  standard  to  LEGACY  SOFTWARE. Some geographies may require the MANUFACTURER to show conformity to the standard to obtain regulatory approval of  the MEDICAL  DEVICE SOFTWARE, even if that software was designed prior   to the existence of the current version of the standard (LEGACY SOFTWARE). In this case, the requirements in 4.4 provide a method for the the MANUFACTURER to demonstrate compliance of LEGACY SOFTWARE to the  standard.

A MANUFACTURER may determine that retrospective documentation of an already finished development-lifecycle performed as an isolated activity does not result in the reduction of RISK associated with the use of the product. The process results in the identification of a subset of

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 45 –

+AMD1:2015 CSV  IEC 2015

ACTIVITIES defined in this standard which does  result in  reduction of  RISK. Some additional  goals implicit in the process are:

- required ACTIVITIES and resulting documentation should rely on and make use of, wherever possible, existing documentation, and

- a MANUFACTURER should utilize resources as effectively as possible to effect a reduction of

RISK.

In addition to a plan identifying the subset of ACTIVITIES to execute, the process also results in objective evidence supporting safe continued use of the LEGACY SOFTWARE and a summary rationale for this conclusion.

The RISKS associated with the planned continued use of the LEGACY SOFTWARE depend on the context in which the LEGACY SOFTWARE will be used to create a SOFTWARE SYSTEM. The MANUFACTURER will document all identified MEDICAL DEVICE HAZARDS associated with the LEGACY SOFTWARE.

Subclause 4.4 requires a comprehensive assessment of available post-production field data obtained for the LEGACY SOFTWARE during the time it has been in production and use. Typical sources of post-production data include:

- adverse events attributable to the device,

- feedback received from users of the device,  and

- ANOMALIES  discovered by the MANUFACTURER.

Though no consensus exists for a method of prospectively estimating quantitatively the probability of occurrence of a software failure, such information may be available for LEGACY SOFTWARE, based on the usage of such software and EVALUATION  of post-production data. If it    is possible in such cases to quantitatively estimate the probability of events in the sequence, a quantitative value may be used for expressing the probability of the entire sequence of events occurring. If such quantitative estimation is not possible, considering a worst case probability is appropriate, and the probability for the software failure occurring should be assumed to be  1.

The MANUFACTURER determination of how the LEGACY SOFTWARE will be used in the  overall MEDICAL DEVICE SYSTEM ARCHITECTURE is input to the assessment of RISK. The RISKS to be considered vary accordingly.

- When LEGACY SOFTWARE has been safely and reliably used and the MANUFACTURER  wishes  to continue use of the LEGACY SOFTWARE, the rationale for continued use rests primarily on the assessment of RISK based on post-production  records.

- When LEGACY SOFTWARE is reused to create a new SOFTWARE SYSTEM, the intended use of the LEGACY SOFTWARE might be different from its original intended use. In this case the RISK assessment must take into account the modified set of HAZARDOUS SITUATIONS which can arise due to failures of the LEGACY  SOFTWARE.

- A reused LEGACY SOFTWARE may be used for similar intended use but integrated into a new SOFTWARE SYSTEM.  In this case the RISK  assessment should take into account modification of architectural RISK CONTROL measures according to  5.3.

When LEGACY SOFTWARE will be changed and used within a new SOFTWARE SYSTEM, the MANUFACTURER should consider how the existing records of safe and reliable operation may be invalidated by the changes.

Changes to the LEGACY SOFTWARE should be performed according to Clauses 4 to 9 of this standard, including assessment of impact to RISK CONTROL measures according to 7.4. In the case of LEGACY SOFTWARE, existing RISK CONTROL measures may not be fully documented and special care should be taken to EVALUATE the potential impact of changes, utilizing available documented design records as well as expertise of individuals having knowledge of the   system.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 46 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

According to 4.4, the MANUFACTURER performs a gap analysis in order to determine  the  available documentation including objective evidence of performed TASKS done during development of the LEGACY SOFTWARE  and compared to 5.2, 5.3, 5.7, and Clause 7. Typical  steps to accomplish this gap analysis  include

- identification of the LEGACY SOFTWARE, including VERSION, revision and any other means, required for clear identification;

- EVALUATION of existing DELIVERABLES corresponding to the deliverables required by 5.2, 5.3, 5.7, and Clause 7;

- EVALUATION of available objective evidence, documenting the previously applied software development lifecycle model (as appropriate);

- EVALUATION of the adequacy of existing RISK MANAGEMENT documentation, taking ISO 14971 into account.

Taking the performed gap analysis into account, the MANUFACTURER will EVALUATE the potential reduction in RISK resulting from the generation of the missing DELIVERABLES and associated ACTIVITIES, and create a plan to perform ACTIVITIES and generate DELIVERABLES to close these gaps.

Reduction of RISK should balance the benefit of applying the software development process according to Clause 5 against the possibility that modification of the LEGACY SOFTWARE without  full knowledge of its development history could introduce new defects that increase the  risk. Some of the elements of Clause 5 may be assessed to have little to no reduction of RISK when done after the fact. For example, detailed design and unit verification reduce RISK primarily during the process of developing new software or refactoring existing software. If these  objectives are not planned, performing the ACTIVITIES in isolation may create documentation but lead to no reduction in RISK.

At a minimum, the gap closure plan addresses missing SOFTWARE SYSTEM test records. If these  do not exist or are not suitable to support a rationale to continue use of the LEGACY SOFTWARE, the gap closure plan should include creation of SOFTWARE SYSTEM requirements at a functional level according to 5.2 and tests according to  5.7.

The documented rationale for continued use of the LEGACY SOFTWARE builds on the available objective evidence and analysis obtained in the course of  assessing the  RISK  and  creating a gap closure plan appropriate for the context of LEGACY SOFTARE   reuse.

The rationale makes a positive case for the safe and reliable performace of the LEGACY  SOFTWARE in the planned reuse context, taking into account both the post-production records available for the LEGACY SOFTWARE and the RISK CONTROL MEASURES affected by filling process gaps.

After LEGACY SOFTWARE has been re-used according to 4.4, those parts  of  the  LEGACY  SOFTWARE for which gaps in DELIVERABLES  remain, continue to be LEGACY  SOFTWARE  and may  be considered for further re-use again according to 4.4. When gaps in deliverables are closed   by changing the LEGACY  SOFTWARE, the changes should be performed according to Clauses 4     to 9 of this standard.

  - Software development PROCESS

#### Software development planning

The objective of this ACTIVITY is to plan the software development  TASKS  to  reduce  RISKS  caused by software, communicate procedures and goals to members of the development team, and ensure that SYSTEM quality requirements for the MEDICAL DEVICE SOFTWARE are   met.

The software development planning ACTIVITY can document TASKS in a single plan or in multiple plans. Some MANUFACTURERS  might have established policies and procedures that apply to    the

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 47 –

+AMD1:2015 CSV  IEC 2015

development of all their MEDICAL DEVICE SOFTWARE. In this case the plan can simply reference the existing policies and procedures. Some MANUFACTURERS might prepare a  plan or  set of plans specific to the development of each MEDICAL DEVICE SOFTWARE that spell out in detail specific ACTIVITIES and reference general procedures. Another possibility is that a plan or set of plans is tailored for the development of each MEDICAL  DEVICE  SOFTWARE. The planning should   be specified at the level of detail necessary to carry out the development PROCESS  and should   be proportional to the RISK. For example, SYSTEMS or items with higher RISK would be subject to    a development PROCESS with more rigor and TASKS should be spelled out in greater   detail.

Planning is an iterative ACTIVITY that should be re-examined and updated as development progresses. The plan can evolve to incorporate more and better information as more is understood about the SYSTEM and the level of effort needed to develop the  SYSTEM.  For  example, a SYSTEM’s initial software safety classification can change as a result of exercising  the RISK MANAGEMENT PROCESS and development of the software ARCHITECTURE. Or it might be decided that a SOUP be incorporated into the SYSTEM. It is important that the plan(s) be updated to reflect current knowledge of  the SYSTEM  and the level of  rigor needed for the SYSTEM  or items in the SYSTEM to enable proper control over the development   PROCESS.

#### Software requirements analysis

This ACTIVITY requires the MANUFACTURER to establish and verify the software requirements for the MEDICAL DEVICE SOFTWARE. Establishing verifiable requirements is essential for determining what is to be built, for determining that the MEDICAL DEVICE SOFTWARE exhibits acceptable behaviour, and for demonstrating that the completed  MEDICAL  DEVICE  SOFTWARE  is  ready for use. To demonstrate that the requirements have been implemented as desired, each  requirement should be stated in such a way that objective criteria can be established to  determine whether it has been implemented correctly. If the device RISK MANAGEMENT PROCESS imposes requirements on the software to control identified RISKS, these requirements are to be identified in the software requirements in such a way as to make it possible to trace the RISK CONTROL measures to the software requirements. All software requirements should  be  identified in such a way as to make it possible to demonstrate TRACEABILITY between the requirement and SOFTWARE SYSTEM testing. If regulatory approval in some countries requires conformance to specific regulations or international standards, this conformance requirement should be documented in the software requirements. Because the software requirements establish what is to be implemented in the software, an evaluation of the  requirements  is required before the requirements analysis ACTIVITY is  complete.

An area of frequent confusion is the distinction between customer needs, design  inputs,  software requirements, software functional specifications, and software design specifications. Design inputs are the interpretation of customer needs into formally  documented  MEDICAL  DEVICE requirements. Software requirements are the formally  documented  specifications  of what the software does to meet the customer needs and the design inputs. Software functional specifications are often included with the software requirements and define in detail what the software does to meet its requirements even  though  many different alternatives  might  also meet the requirements. Software design specifications define how the software will  be  designed and decomposed to implement its requirements and functional  specifications.

Traditionally, software requirements, functional specifications, and design specifications have been written as a set of one or more documents. It is now feasible to consider this information   as data items within a common database. Each item would have one or more attributes that would define its purpose and linkage to other items in the database. This approach allows presentation and printing of different views of the information best suited for each  set  of intended users (e.g., marketing, MANUFACTURERS, testers, auditors) and supports TRACEABILITY  to demonstrate adequate implementation and the extent to which test cases test the requirements. Tools to support this approach can be as simple as a hypertext document using HTML hyperlinks or as complex and capable as computer aided software engineering (CASE) tools and requirements analysis tools.

The SYSTEM requirements PROCESS is out of scope of this standard. However, the decision to implement MEDICAL DEVICE functionality with software is normally made during SYSTEM design. Some  or  all  of  the  SYSTEM  requirements  are  allocated  to  be  implemented in  software. The

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 48 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

software requirements analysis ACTIVITY consists of analyzing the requirements allocated to software by the SYSTEM requirements PROCESS and deriving a comprehensive set of software requirements that reflect the allocated requirements.

To ensure the integrity of the SYSTEM, the MANUFACTURER should provide a mechanism for negotiating changes and clarifications to the SYSTEM requirements to correct impracticalities, inconsistencies or ambiguities in either the parent SYSTEM requirements or the software requirements.

The PROCESS of capture and analysis of SYSTEM and software requirements can be  iterative.  This standard does not intend to require the PROCESSES to  be  rigidly  segregated  into  two layers. In practice, SYSTEM ARCHITECTURE and software ARCHITECTURE are often outlined simultaneously and the SYSTEM and software requirements are subsequently documented in a layered form.

    - Software ARCHITECTURAL design

This ACTIVITY requires the MANUFACTURER to define the major structural components of the software and identify their key responsibilities, their externally visible properties, and the relationship among them. If the behaviour of a component can affect other components, that behavior should be described in the software ARCHITECTURE. This description is especially important for behaviour that can affect components of the MEDICAL DEVICE that are outside the software (see 5.3.5 and B.4.3). ARCHITECTURAL decisions are extremely important for implementing RISK CONTROL measures. Without understanding (and documenting) the behaviour of a  component that can affect other components, it  will be  nearly impossible to  show that the SYSTEM is safe. A software ARCHITECTURE is necessary to ensure the correct implementation of the software requirements. The software ARCHITECTURE is not  complete  unless all software requirements can be implemented by the identified SOFTWARE  ITEMS.  Because the design and implementation of the software is dependent on the ARCHITECTURE, the ARCHITECTURE is VERIFIED to complete this ACTIVITY. VERIFICATION of the ARCHITECTURE  is generally done by technical EVALUATION.

The software safety classification of SOFTWARE ITEMS during the  software  ARCHITECTURE  ACTIVITY creates a basis for the subsequent choice of software PROCESSES. The records of classification are placed under change control as part of the RISK MANAGEMENT   FILE.

Many subsequent events might invalidate the classification. These include, for  example:

- changes of SYSTEM specification, software specification or  ARCHITECTURE;

- discovery of errors in the RISK ANALYSIS, especially unforeseen HAZARDS;  and

- discovery of the infeasibility of a requirement, especially a RISK CONTROL   measure;

Therefore, during all ACTIVITIES following the design of the software ARCHITECTURE, the classification of the SOFTWARE SYSTEM and SOFTWARE ITEMS should be re-EVALUATED and might need to be revised. This would trigger rework to apply additional PROCESSES  to a  SOFTWARE   ITEM as a result of its upgrading to a higher class. The software configuration management PROCESS (Clause 8) is used to ensure that all necessary rework is identified and   completed.

#### Software detailed design

This ACTIVITY requires the MANUFACTURER to refine the SOFTWARE ITEMS  and interfaces defined  in the ARCHITECTURE to create SOFTWARE UNITS  and their interfaces. Although SOFTWARE UNITS  are often thought of as being a single function or module, this view is not always appropriate. This standard has defined SOFTWARE UNIT to be a SOFTWARE ITEM that is not subdivided into smaller items. SOFTWARE UNITS can be tested separately. The MANUFACTURER should define the level of detail of the SOFTWARE UNIT. Detailed design specifies algorithms, data representations, interfaces among different SOFTWARE UNITS, and interfaces between SOFTWARE UNITS and data structures. Detailed design must also be concerned with the packaging of the SOFTWARE PRODUCT. It is necessary to define the design of the SOFTWARE UNITS and the interfaces in sufficient detail to permit its SAFETY and effectiveness to be objectively VERIFIED where this can be ensured using other requirements or design documentation. It should be complete enough

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 49 –

+AMD1:2015 CSV  IEC 2015

that the programmer is not required to  make ad  hoc  design decisions. Detailed design must also be concerned with the architecture of the MEDICAL DEVICE   SOFTWARE.

A SOFTWARE ITEM can be decomposed so that only a few of  the  new  SOFTWARE  ITEMS  implement the SAFETY-related requirement of the original SOFTWARE ITEM. The remaining SOFTWARE ITEMS do not implement SAFETY-related functions and  can  be  reclassified  into  a lower software safety class. However, the decision to do this is in itself part of the RISK MANAGEMENT PROCESS, and is documented in the RISK MANAGEMENT   FILE.

Because implementation depends on detailed design, it is necessary to verify the  detailed  design before the ACTIVITY is complete. VERIFICATION of detailed design is generally done by a technical EVALUATION. Subclause 5.4.4 requires the MANUFACTURER to verify the outputs of the detailed design ACTIVITIES. The design specifies how the requirements are to be implemented. VERIFICATION of the design provides assurance that it implements the software ARCHITECTURE  and is free from contradiction with the software  ARCHITECTURE.

If the design contains defects, the code will not implement the requirements  correctly.

When present in the design, the MANUFACTURER should verify design characteristics which the

MANUFACTURER believes are important for SAFETY. Examples of these characteristics  include:

- implementation of the intended events, inputs, outputs, interfaces, logic flow, allocation of CPU, allocation of memory resources, error and exception definition, error and exception isolation, and error recovery;

- definition of the default state, in which all faults that can result in a hazardous situation are addressed, with events and transitions;

- initialization of variables, memory management; and

- cold and warm resets, standby, and other state changes that can affect the RISK CONTROL

measures.

    - SOFTWARE UNIT implementation and verification

This ACTIVITY requires the MANUFACTURER  to write and verify the code for the SOFTWARE UNITS.  The detailed design is to be translated into source code. Coding represents the point where decomposition of  the specifications ends and composition of the executable software begins.  To consistently achieve the desirable code characteristics, coding standards should be used to specify a preferred coding style. Examples of coding standards include requirements for understandability, language usage rules or restrictions, and complexity management. The code for each unit is VERIFIED to ensure that it functions as specified by the detailed design and that     it complies with the specified coding standards.

Subclause 5.5.5 requires the MANUFACTURER to verify the code. If the code does not implement  the design correctly, the MEDICAL DEVICE SOFTWARE will not perform as   intended.

#### Software integration and integration testing

This ACTIVITY  requires the MANUFACTURER  to plan and execute integration of  SOFTWARE  UNITS   into aggregate SOFTWARE ITEMS as well as integration of SOFTWARE ITEMS  into  higher  aggregated SOFTWARE ITEMS and to verify that the resulting SOFTWARE ITEMS behave  as  intended.

The approach to integration can range from non-incremental integration to any form of incremental integration. The properties of the SOFTWARE ITEM being assembled dictate  the chosen method of integration.

Software integration testing focuses on the transfer of data and control across a  SOFTWARE ITEM’s internal and external interfaces. External interfaces are those with other software, including operating system software, and MEDICAL DEVICE  hardware.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 50 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

The rigor of integration testing and the level of detail of the documentation associated with integration testing should be commensurate with the RISK associated with the device, the device’s dependence on software for potentially hazardous functions, and the role of specific SOFTWARE ITEMS in higher RISK device functions. For example, although all SOFTWARE ITEMS should be tested, items that have an effect on SAFETY should be subject to more  direct,  thorough, and detailed tests.

As applicable, integration testing demonstrates program behaviour  at  the  boundaries  of  its input and output domains and confirms program responses to invalid, unexpected, and special inputs. The program’s actions are revealed when given combinations of inputs or unexpected sequences of inputs, or when defined timing requirements are violated. The test requirements    in the plan should include, as  appropriate, the types of  white box testing to be performed as part of integration testing.

White box testing, also known as glass box, structural, clear box and open box testing, is a testing technique where explicit knowledge of the internal workings of the SOFTWARE ITEM being tested are used to select the test data. White box testing uses specific knowledge of the SOFTWARE ITEM to examine outputs. The test is accurate only if the tester knows what the SOFTWARE ITEM  is supposed to do. The tester can then see if the SOFTWARE ITEM diverges from  its intended goal. White box testing cannot guarantee that the complete specification has been implemented since it is focused on testing the implementation of the SOFTWARE ITEM. Black box testing, also known as behavioural, functional, opaque-box, and closed-box testing, is focused  on testing the functional specification and it cannot guarantee that all parts of the implementation have been tested. Thus black box testing is  testing against the specification  and will discover faults of omission, indicating that part of the specification has not  been  fulfilled.  White  box  testing  is testing  against  the  implementation  and   will   discover  faults of commission, indicating that part of the implementation is faulty. In order to fully test MEDICAL DEVICE SOFTWARE both black and white box testing might be   required.

The plans and test documentation identified in 5.6 and 5.7 can be individual documents tied to specific phases of development or evolutionary prototypes. They also might be combined so a single document or set of documents covers the requirements of multiple subsections. All or portions of the documents could be incorporated into higher level project documents such as a software or project quality assurance plan or a comprehensive test plan that addresses all aspects of testing for hardware and software. In these cases, a cross reference should be created that identifies how the various project documents relate to each of the software integration TASKS.

Software integration testing can be performed in a simulated environment, on actual target hardware, or on the full MEDICAL  DEVICE.

Subclause 5.6.2 requires the MANUFACTURER to verify the output of the software integration ACTIVITY. The output of the software integration ACTIVITY is the integrated SOFTWARE ITEMS. These integrated SOFTWARE ITEMS must function properly for the entire MEDICAL  DEVICE  SOFTWARE to function correctly and safely.

    - SOFTWARE SYSTEM testing

This ACTIVITY requires the MANUFACTURER  to verify the software’s functionality by verifying that  the requirements for the software have been successfully  implemented.

SOFTWARE SYSTEM testing demonstrates that the specified functionality exists. This testing VERIFIES the functionality and performance of the program as built with respect to the requirements for the software.

SOFTWARE SYSTEM testing focuses on functional (black box) testing, although it might be  desirable to use white box (see previous section) methods to  more  efficiently  accomplish certain tests, initiate stress conditions or faults, or increase code coverage of the qualification tests.  The  organization  of   testing  by  types   and  test  stage  is   flexible,  but  coverage     of

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 51 –

+AMD1:2015 CSV  IEC 2015

requirements, RISK CONTROL, usability, and test types (e.g., fault, installation, stress) should be demonstrated and documented.

SOFTWARE SYSTEM testing tests the integrated software and can be performed in a simulated environment, on actual target hardware, or on the full MEDICAL  DEVICE.

When a change is made to a SOFTWARE SYSTEM (even a small change), the degree of REGRESSION TESTING (not just the testing of the individual change) should be determined to ensure that no  unintended side effects have been introduced. This REGRESSION  TESTING  (and the rationale for not fully repeating SOFTWARE SYSTEM testing) should be planned and documented. (See B.6.3).

SOFTWARE SYSTEM test responsibilities can be dispersed, occurring at different locations and being conducted by different organizations. However, regardless of the distribution of TASKS, contractual relations, source of components, or development environment, the device MANUFACTURER retains ultimate responsibility for ensuring that the  software functions properly  for its intended use.

If  ANOMALIES  uncovered during testing can be  repeated, but a  decision has been made not to fix them, then these ANOMALIES need to be EVALUATED in relation to the RISK  analysis to verify  that they do not affect the SAFETY of the device. The root cause and symptoms  of  the  ANOMALIES should be understood, and the rationale for not fixing them should be   documented.

Subclause 5.7.4 requires the results of the SOFTWARE SYSTEM testing be EVALUATED to ensure  that the expected results were obtained.

#### Software release

This ACTIVITY requires the MANUFACTURER to document the VERSION of the MEDICAL DEVICE SOFTWARE being released, specify how it was created, and follow appropriate procedures for release of the software.

The MANUFACTURER should be able to show that the software that was developed using the development PROCESS is  the software that is  being released. The MANUFACTURER  should also  be able to retrieve the software and the tools used for its generation in case it is needed in the future and should store, package, and deliver the software in a manner that minimizes the software from being damaged or misused. Defined procedures should be established to ensure that these TASKS are performed appropriately and with consistent  results.

  - Software maintenance PROCESS

#### Establish software maintenance plan

The software maintenance PROCESS differs from the software development  PROCESS  in  two ways:

- The MANUFACTURER is permitted to use a smaller PROCESS than the full  software  development PROCESS to implement rapid changes in response to urgent  problems.

- In responding to software PROBLEMS REPORTS relating to released product,  the MANUFACTURER not only addresses the problem but also satisfies local regulations (typically  by running a pro-active surveillance scheme for collecting problem data from the field and communicating with users and regulators about the  problem).

Subclause 6.1 requires these PROCESSES to be established in a maintenance   plan.

This ACTIVITY requires the MANUFACTURER to create or identify procedures for implementing maintenance ACTIVITIES and TASKS. To implement corrective actions, control changes during maintenance, and manage release of revised software, the  MANUFACTURER  should document and resolve reported problems and requests from users, as well as manage modifications to

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 52 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

the MEDICAL DEVICE SOFTWARE. This PROCESS is activated when the MEDICAL DEVICE SOFTWARE undergoes modifications to code and associated documentation because of either a problem or the need for improvement or adaptation. The objective is to modify released MEDICAL DEVICE SOFTWARE while preserving its integrity. This PROCESS  includes  migration  of  the  MEDICAL DEVICE SOFTWARE to environments or platforms for which it was not originally released. The ACTIVITIES provided in this clause are specific to the maintenance PROCESS; however, the maintenance PROCESS might use other PROCESSES in this  standard.

The MANUFACTURER  needs  to plan how the ACTIVITIES  and TASKS  of the maintenance    PROCESS

will be performed.

#### Problem and modification analysis

This ACTIVITY requires the MANUFACTURER to analyze feedback for its effect; verify reported problems; and consider, select, and obtain approval for implementing a modification option. Problems and other requests for changes can affect the performance, SAFETY, or regulatory clearance of a MEDICAL DEVICE. An analysis is necessary to determine whether any effects exist because of a PROBLEM REPORT or whether any effects will result from a modification to correct a problem or implement a request. It is especially important to verify through trace or regression analysis that the RISK CONTROL measures built into the device are not adversely changed or modified by the software change that is being implemented as part  of the  software  maintenance ACTIVITY. It is also important to verify that the modified software does not cause a HAZARDOUS SITUATION or mitigate a RISK in software that previously did not cause a HAZARDOUS SITUATION or mitigate RISKS. The software safety classification of a SOFTWARE ITEM might have changed if the software modification now can cause a HAZARD or mitigate a   RISK.

It is important to distinguish between software maintenance (Clause 6) and software problem resolution (Clause 9).

The focus of the software maintenance PROCESS is an adequate response to feedback arising after release of the MEDICAL DEVICE SOFTWARE. As part of a MEDICAL DEVICE, the software maintenance PROCESS needs to ensure  that:

- SAFETY-related PROBLEM REPORTS are addressed and reported to appropriate regulatory authorities and affected users;

- MEDICAL DEVICE SOFTWARE is re-validated and re-released after modification with formal controls that ensure the rectification of the problem and the avoidance of further   problems;

- the MANUFACTURER considers what other MEDICAL DEVICE SOFTWARE might be affected and takes appropriate action.

The focus of software problem resolution is the operation of a comprehensive control system that:

- analyses PROBLEM REPORTS and identifies all the implications of the   problem;

- decides on a number of changes and identifies all their  side-effects;

- implements the changes while maintaining the consistency of the software CONFIGURATION ITEMS  including the RISK MANAGEMENT FILE;

- VERIFIES the implementation of the  changes.

The software maintenance PROCESS uses the software problem resolution PROCESS. The  software maintenance PROCESS handles the high-level decisions about the PROBLEM REPORT (whether a problem exists, whether it has a significant effect on SAFETY, what changes are needed and when to implement them), and uses the software problem resolution PROCESS to analyse the PROBLEM REPORT to discover all the implications and to generate possible CHANGE REQUESTS which identify all the CONFIGURATION ITEMS that need to be changed and all the VERIFICATION steps that are  necessary.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 53 –

+AMD1:2015 CSV  IEC 2015

#### Modification implementation

This ACTIVITY requires that the MANUFACTURER use an established PROCESS to make the modification. If a maintenance PROCESS has not been defined, the appropriate development PROCESS TASKS can be used to make the modification. The MANUFACTURER should also ensure that the modification does not cause a negative effect on other parts of the MEDICAL DEVICE SOFTWARE. Unless the MEDICAL DEVICE SOFTWARE is treated as a new development, analysis of the effect of a modification on the entire MEDICAL DEVICE SOFTWARE is necessary. Regression analysis and testing are employed to provide assurance that a change has  not  created problems elsewhere in the MEDICAL DEVICE SOFTWARE. Regression analysis is the determination of the impact of a change based on review of the relevant documentation (e.g., software requirements specification, software design specification, source code, test plans, test cases, test scripts, etc.) in order to identify the necessary regression tests to  be  run.  Regression testing is the rerunning of test cases that a program has previously executed correctly and comparing the current result to the previous result in order to detect unintended effects of a software change. A rationale must be made that justifies the amount of  REGRESSION  TESTING  that will be performed to ensure that the portions of the MEDICAL DEVICE SOFTWARE not being modified still perform as they did before the modification was  made.

B.7     Software RISK MANAGEMENT  PROCESS

Software RISK MANAGEMENT is a part of overall MEDICAL DEVICE RISK MANAGEMENT and cannot be adequately addressed in isolation. This standard requires the use of a RISK MANAGEMENT  PROCESS that is compliant with ISO 14971. RISK MANAGEMENT as defined in ISO 14971 deals specifically with a framework for effective management of the RISKS associated with the use of MEDICAL DEVICES. One portion of ISO 14971 pertains to  control of  identified RISKS  associated with each HAZARD identified during the RISK ANALYSIS. The software RISK  MANAGEMENT  PROCESS  in this standard is intended to provide additional requirements for RISK CONTROL for software, including software that has been identified during the RISK  ANALYSIS  as potentially contributing  to a hazardous situation, or software that is used to control MEDICAL DEVICE RISKS. The software RISK MANAGEMENT PROCESS is included in this standard for two   reasons.

- the intended audience of this standard needs to understand minimum requirements for RISK CONTROL measures in their area of  responsibility—software;

- the general RISK MANAGEMENT standard, ISO 14971, provided as  a  normative reference in this standard, does not specifically address the RISK CONTROL of software  and  the  placement of RISK CONTROL in the software development life  cycle.

Software RISK MANAGEMENT is a part of overall MEDICAL DEVICE RISK MANAGEMENT. Plans, procedures, and documentation required for the software RISK MANAGEMENT ACTIVITIES can be a series of separate documents or a single document, or they can be integrated with the MEDICAL DEVICE RISK MANAGEMENT ACTIVITIES and documentation as long as all requirements in this standard are met.

#### B.7.1       Analysis of software contributing to hazardous situations

It is expected that the device HAZARD analysis will identify hazardous situations  and corresponding RISK CONTROL measures to reduce the probability and/or severity of those hazardous situations to an acceptable level. It is also expected  that  the  RISK  CONTROL  measures will be assigned to software functions that are expected to implement those RISK CONTROL measures.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 54 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

However, it is not expected that all device hazardous situations can be identified until the software ARCHITECTURE has been produced. At that time it is known how software functions will  be implemented in software components, and the practicality of the RISK CONTROL measures assigned to software functions can be EVALUATED. At that time the  device  HAZARD  analysis should be revised to include:

- revised hazardous situations;

- revised RISK CONTROL measures and software  requirements;

- new hazardous situations arising from  software, for example hazardous situations related    to human factors.

The software ARCHITECTURE should include credible strategies for segregating software components so that they do not interact in unsafe  ways.

### Software configuration management PROCESS

The software configuration management PROCESS is a PROCESS of applying administrative and technical procedures throughout the software life cycle to identify and define SOFTWARE ITEMS, including documentation, in a SYSTEM; control modifications and releases of the items; and document and report the status of the items and CHANGE REQUESTS. Software configuration management is necessary to recreate a SOFTWARE ITEM, to identify its constituent parts, and to provide a history of the changes that have been made to  it.

#### Configuration identification

This ACTIVITY requires the MANUFACTURER to uniquely identify software CONFIGURATION ITEMS and their VERSIONS. This identification is necessary to  identify the software CONFIGURATION  ITEMS  and the VERSIONS that are included in the MEDICAL DEVICE   SOFTWARE.

#### Change control

This ACTIVITY requires the MANUFACTURER to control changes of the software CONFIGURATION ITEMS and to document information identifying CHANGE REQUESTS and providing documentation about their disposition. This ACTIVITY is necessary to ensure that unauthorized or unintended changes are not made to the software CONFIGURATION ITEMS and to ensure  that  approved CHANGE REQUESTS are implemented fully and  verified.

CHANGE REQUESTS can be approved by a change control board or by a  manager or technical lead according to the software configuration management plan.  Approved  CHANGE  REQUESTS are made traceable to the actual modification and VERIFICATION of the  software.  The requirement is that each actual change be linked to a CHANGE REQUEST and that documentation exists to show that the CHANGE REQUEST was approved. The documentation might be change control board minutes, an approval signature, or a record in a  database.

#### Configuration status accounting

This ACTIVITY requires the MANUFACTURER to maintain records of the history of the software CONFIGURATION ITEMS. This ACTIVITY is necessary to determine when and why changes were made. Access to this information is necessary to ensure that software CONFIGURATION ITEMS contain only authorized modifications.

### B.9     Software problem resolution PROCESS

The software problem resolution PROCESS is a PROCESS for analyzing and resolving  the  problems (including non-conformances), whatever their nature or source, including those discovered during the execution of development, maintenance, or other PROCESSES. The objective is to provide a timely, responsible, and documented means to ensure that discovered problems  are  analyzed  and  resolved  and  that  trends  are  recognized.  This        PROCESS   is

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 55 –

+AMD1:2015 CSV  IEC 2015

sometimes called “defect tracking” in software engineering literature. It is called “problem resolution” in ISO/IEC 12207 [9] and IEC 60601-1-4 [2], Amendment 1. We have chosen to call    it “software problem resolution” in this  standard.

This ACTIVITY requires that the MANUFACTURER use the  software  problem  resolution  PROCESS when a problem or non-conformance is identified. This ACTIVITY is necessary to ensure that discovered problems are analyzed and EVALUATED for possible relevance to  SAFETY  (as specified in ISO 14971).

Software development plan(s) or procedures, as required in 5.1, are to address how problems    or  non-conformances will be handled. This includes specifying at each stage of the life cycle    the aspects of the software problem resolution PROCESS that will be formal and documented as well as when problems and nonconformities are to be entered into the software problem resolution PROCESS.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 56 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

# Annex C

## (informative)

Relationship to other standards

### General

This standard applies to the development and maintenance of MEDICAL DEVICE SOFTWARE. The software is considered a subsystem of the MEDICAL DEVICE or is itself a MEDICAL DEVICE. This standard is to be used together with other appropriate standards when developing a MEDICAL DEVICE.

MEDICAL DEVICE management standards such as ISO 13485 [8] (see C.2 and Annex D) and ISO 14971 (see Annex C.3) provide a management environment that lays a foundation for an organization to develop products. Safety standards such as  IEC 60601-1 [1] (see Annex C.4)  and IEC 61010-1 [5] (see Annex C.5) give specific direction for creating safe MEDICAL DEVICES. When software is a part of these MEDICAL DEVICES, IEC 62304 provides more detailed direction on what is required to develop and maintain safe MEDICAL DEVICE SOFTWARE. Many other standards such as ISO/IEC 12207 [9] (see Annex C.6), IEC 61508-3 [4] (see Annex C.7) and ISO/IEC 90003 [15] can be looked to as a source of methods, tools and techniques that can be used to implement the requirements in IEC 62304. Figure C.1 shows the relationship of these standards.

Medical device product standards

IEC 60601-1

IEC 61010-1

IEC 82304-1

Medical device management standards ISO 14971

ISO 13485

Gives specific direction for creation of a safe medical device

requires

Where clauses or requirements from other standards are quoted, defined terms in the quoted items are terms that are defined in the other standard, not defined terms in this   standard.

Lays out a foundation to

develop a medical device

affects

affects

Medical device process standards

IEC 62304

IEC 62366-1

affects

Implementation of medical device software

inspires

Other sources of information IEC/ISO 12207

IEC 61508-3

IEC/ISO 90003, ...

Gives additional guidelines, techniques, etc that may be used

Gives detailed direction

how to develop and maintain safe software system

IEC

Figure C.1 – Relationship of key MEDICAL DEVICE standards to IEC  62304

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 57 –

+AMD1:2015 CSV  IEC 2015

### Relationship to ISO 13485

This standard requires that the MANUFACTURER  employs a quality management system.  When   a MANUFACTURER uses ISO 13485 [8], the requirements of ISO 62304 directly relate to some of the requirements of ISO 13485 as shown in Table  C.1.

#### Table C.1 – Relationship to ISO 13485:2003

| IEC 62304 clause | Related clause of ISO 13485:2003 |
| --- | --- |
| 5.1 Software development planning | 7.3.1 Design and development planning |
| 5.2 Software requirements analysis | 7.3.2 Design and development inputs |
| 5.3 Software ARCHITECTURAL design |  |
| 5.4 Software detailed design |  |
| 5.5 SOFTWARE UNIT implementation and verification |  |
| 5.6 Software integration and integration testing |  |
| 5.7 SOFTWARE SYSTEM testing | - Design and development outputs - Design and development review |
| 5.8 Software release | - Design and development verification - Design and development validation |
| 6.1 Establish software maintenance plan | 7.3.7 Control of design and development changes |
| 6.2 Problem and modification analysis |  |
| 6.3 Modification implementation | - Design and development verification - Design and development validation |
| 7.1 Analysis of software contributing to hazardous situations |  |
| 7.2 RISK CONTROL measures |  |
| 7.3 VERIFICATION of RISK CONTROL measures |  |
| 7.4 RISK MANAGEMENT of software changes |  |
| 8.1 Configuration identification | 7.5.3 Identification and TRACEABILITY |
| 8.2 Change control | 7.5.3 Identification and TRACEABILITY |
| 8.3 Configuration status accounting |  |
| 9 Software problem resolution PROCESS |  |

  - Relationship to ISO 14971

Table C.2 shows the areas where IEC 62304 amplifies requirements for the RISK MANAGEMENT PROCESS required by ISO 14971.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 58 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

#### Table C.2 – Relationship to ISO 14971:2007

| ISO 14971:2007 clause | Related clause of IEC 62304 |  |
| --- | --- | --- |
| 4.1 RISK ANALYSIS process |  |  |
| 4.2 Intended use and identification of characteristics related to the SAFETY of the MEDICAL DEVICE |  |  |
| 4.3 Identification of HAZARDS | 7.1 | Analysis of software contributing to HAZARDOUS SITUATIONS |
| 4.4 Estimation of the RISK(S) for each HAZARDOUS SITUATION | 4.3 | Software safety classification |
| 5 RISK evaluation |  |  |
| 6.1 RISK reduction |  |  |
| 6.2 RISK CONTROL option analysis | 7.2.1 | Define RISK CONTROL measures |
| 6.3 Implementation of RISK CONTROL measures | 7.2.2 7.3.1 | RISK CONTROL measures implemented in software Verify RISK CONTROL measures |
| 6.4 RESIDUAL RISK evaluation |  |  |
| 6.5 RISK/benefit analysis |  |  |
| 6.6 RISKS arising from RISK CONTROL MEASURES | 7.3.2 | Document any new sequences of events |
| 6.7 Completeness of RISK CONTROL |  |  |
| 7 Evaluation of overall RESIDUAL RISK acceptability |  |  |
| 8 RISK MANAGEMENT report | 7.3.3 | Document TRACEABILITY |
| 9 Production and post-production information | 7.4 | RISK MANAGEMENT of software changes |

  - Relationship to PEMS requirements of IEC  60601-1:2005

+  IEC 606011:2005/AMD1:2012

    - General

Requirements for software are a subset of the requirements for a programmable electrical medical system (PEMS). This standard identifies requirements for software which are  in addition to, but not incompatible with, the requirements of IEC 60601-1:2005 + IEC 60601-  1:2005 /AMD1:2012 [1] for PEMS. Because PEMS include elements that are not software, not  all of the requirements of IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 for PEMS are addressed  in this  standard.  With  the  publication  of IEC 60601-1:2005 + IEC   60601-   1:2005 /AMD1:2012, IEC 62304 is now a normative reference of IEC 60601-1 and compliance with Clause 14 of IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 (and  thus  compliance  with  the  standard)  requires  compliance  with  parts  of IEC 62304  (not  with  the  whole  of IEC 62304 because IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 does not require compliance with post-production and maintenance requirements of IEC 62304). Finally, it is important to remember that IEC 60601-1:2005 + IEC 60601-1:2005/AMD1:2012 is only used if the software is part of a PEMS and not if the software is itself a MEDICAL DEVICE.

#### Software relationship to PEMS development

By using the V-model illustrated in Figure C.2 to describe what occurs during a PEMS development, it can be seen that the requirements of this software standard apply at the PEMS component level, from the specification of the software requirements to the integration of the SOFTWARE ITEMS into a SOFTWARE SYSTEM. This SOFTWARE SYSTEM is a part of a programmable electrical subsystem (PESS), which is a part of a  PEMS.

IEC 62304:2006	– 59 –

+AMD1:2015 CSV  IEC 2015

| Software unit VERIFICATION (unit VERIFICATION) | Unit VERIFICATIO |
| --- | --- |
|  | results |

PEMS

requirements capture

PEMS validation

Subsystem (e.g. PESS) integration &

VERIFICATION

Software unit implementation

PEMS

architectural design

PEMS integration &

VERIFICATION

Subsystem (e.g.

PESS)

architectural design

User needs

Validated PEMS

PEMS validation plan

PEMS

validation results

PEMS requirement

specifications

Verified PEMS

PEMS test specification

PEMS VERIFICATION Plan

PEMS

VERIFICATION

results

PEMS architecture specification,

Subsystem (e.g. PESS) requirements  specifications

Verified Subsystem

Subsystem test specification

Subsystem

VERIFICATION

results

Verified software subsystem (component)

Software test specifications

Software  ARCHITECTURE specification

Verified code

N

Software integration & SOFTWARE SYSTEM VERIFICATION

(component integration & verification)

Software architectural design (component design)

Portion of PEMS V-model included in IEC 62304

Software requirements specifications (component  requirements)

Software integration and

VERIFICATION

results

Software detailed design

(unit design)

Key:

Boxes represent typical development lifecycle activities

Solid Arrows indicate typical deliverables transfered into/out of activities Dotted arrows indicate deliverables just to the Risk Management File

Outputs from problem resolution process Inputs to problem resolution process

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC  726/06

#### Figure C.2 – Software as part of the  V-model

    - Development PROCESS

Compliance with the software development PROCESS of this standard (Clause 5) requires that a software development plan be specified and then followed; it does not require  that  any particular life cycle model is  used, but it  does require that the plan include certain ACTIVITIES   and  have  certain  attributes.  These  requirements  relate  to   the   PEMS  requirements  in   IEC 60601-1 for development life cycle, requirement specification, ARCHITECTURE, design and implementation, and VERIFICATION. The requirements in this standard provide greater detail  about software development than those in IEC  60601-1.

    - Maintenance PROCESS

Compliance with the software maintenance PROCESS of this standard (Clause 6) requires that procedures be established and followed when changes to software are made. These require- ments correspond to the requirement in IEC 60601-1 for modification of a PEMS. The requirements in this standard for software maintenance  provide  greater  detail  about  what  must be done for software  maintenance  than  the  requirements  for  PEMS  modification  in IEC 60601-1.

    - Other PROCESSES

The other PROCESSES in this standard specify additional requirements for software beyond the similar requirements for PEMS in IEC 60601-1. In most cases, there is a general requirement for PEMS in IEC 60601-1, which the PROCESSES in this standard expand  upon.

The software RISK MANAGEMENT PROCESS in this standard corresponds to the additional RISK MANAGEMENT requirements identified for PEMS in IEC  60601-1.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 60 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

The software problem resolution PROCESS in this standard corresponds to the  problem  resolution requirement for PEMS in IEC 60601-1.

The software configuration management PROCESS in this standard specifies additional requirements that are not present for PEMS in IEC 60601-1 except for  documentation.

#### Coverage of PEMS requirements in IEC 60601-1:2005

+ IEC 606011:2005 /AMD1:2012

Table C.3 shows the PEMS requirements of IEC 60601-1 and the corresponding requirements  in this standard.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 61 –

+AMD1:2015 CSV  IEC 2015

#### Table C.3 – Relationship to IEC 60601-1

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.1 General The requirements in 14.2 to 14.12 (inclusive) shall apply to PEMS unless: - none of the PROGRAMMABLE ELECTRONIC SUBSYSTEMS (PESS) provides functionality necessary for BASIC SAFETY or ESSENTIAL PERFORMANCE; or - the application of RISK MANAGEMENT as described in 4.2 demonstrates that the failure of the PESS does not lead to an unacceptable RISK. The requirements in 14.13 are applicable to any PEMS intended to be incorporated into an IT- NETWORK whether or not the requirements in 14.2 to 14.12 apply. When the requirements in 14.2 to 14.13 apply, the requirements in subclause 4.3, Clause 5, Clause 7, Clause 8 and Clause 9 of IEC 62304:2006 shall also apply to the development or modification of software for each PESS. | 4.3 Software safety classification The PEMS requirements of IEC 60601-1 would only apply to software safety classes B and C. This standard includes some requirements for software safety class A. The software development PROCESS required for compliance with IEC 60601-1 does not include the post production monitoring and maintenance required by Clause 6 of IEC 62304:2006. |
| 14.2 Documentation The documents required by Clause 14 shall be reviewed, approved, issued and changed in accordance with a formal document control procedure. | 5.1 Software development planning In addition to the specific requirements in the software development planning ACTIVITY, documents that are part of the RISK MANAGEMENT FILE are required to be maintained by ISO 14971. In addition, for documents that are required by the quality system, ISO 13485 [8] requires control of the documents. |
| 14.3 RISK MANAGEMENT PLAN The RISK MANAGEMENT plan required by 4.2.2 shall also include a reference to the PEMS VALIDATION plan (see 14.11). | Not specifically required. There is no specific software validation plan. The PEMS validation plan is at the SYSTEM level and thus is outside the scope of this software standard. This standard does require TRACEABILITY from HAZARD to specific software cause to RISK CONTROL measure to VERIFICATION of the RISK CONTROL measure (see 7.3) |
| 14.4 PEMS DEVELOPMENT LIFE-CYCLE A PEMS DEVELOPMENT LIFE-CYCLE shall be documented. | 5.1 Software development planning 5.1.1 Software development plan The items addressed by the software development plan constitute a SOFTWARE DEVELOPMENT LIFE CYCLE. |
| The PEMS DEVELOPMENT LIFE-CYCLE shall contain a set of defined milestones. |  |
| At each milestone, the ACTIVITIES to be completed and the VERIFICATION methods to be applied to those activities shall be defined. | 5.1.6 Software VERIFICATION planning VERIFICATION TASKS, milestones and acceptance criteria must be planned. |
| Each activity shall be defined including its inputs and outputs. | 5.1.1 Software development plan ACTIVITIES are defined in this standard. Documentation to be produced is defined in each ACTIVITY. |
| Each milestone shall identify the RISK MANAGEMENT ACTIVITIES that must be completed before that milestone. |  |
| The PEMS DEVELOPMENT LIFE-CYCLE shall be tailored for a specific development by making plans which detail ACTIVITIES, milestones and schedules. | 5.1.1 Software development plan This standard allows the development life cycle to be documented in the development plan. This means the development plan contains a tailored development life cycle. |
| The PEMS DEVELOPMENT LIFE-CYCLE shall include documentation requirements. | 5.1.1 Software development plan 5.1.8 Documentation planning |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 62 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.5 Problem resolution Where appropriate, a documented system for problem resolution within and between all phases and ACTIVITIES of the PEMS DEVELOPMENT LIFE- CYCLE shall be developed and maintained. | 9 Software problem resolution PROCESS |
| Depending on the type of product, the problem resolution SYSTEM may: - be documented as a part of the PEMS DEVELOPMENT LIFE-CYCLE; - allow the reporting of potential or existing problems affecting BASIC SAFETY or ESSENTIAL PERFORMANCE; - include an assessment of each problem for associated RISKS; - identify the criteria that must be met for the issue to be closed; - identify the action to be taken to resolve each problem. | 5.1.1 Software development plan 9.1 Prepare PROBLEM REPORTS |
| 14.6 RISK MANAGEMENT PROCESS | 7 Software RISK MANAGEMENT PROCESS |
| 14.6.1 Identification of known and foreseeable HAZARDS When compiling the list of known or foreseeable HAZARDS, the MANUFACTURER shall consider those HAZARDS associated with software and hardware aspects of the PEMS including those associated with the incorporation of the PEMS into an IT-NETWORK, components of third-party origin and legacy subsystems. | 7.1 Analysis of software contributing to HAZARDOUS SITUATIONS This standard does not mention network/data coupling specifically |
| 14.6.2 RISK CONTROL Suitably validated tools and PROCEDURES shall be selected and identified to implement each RISK CONTROL measure. These tools and PROCEDURES shall be appropriate to assure that each RISK CONTROL measure satisfactorily reduces the identified RISK(S). | 5.1.4 Software development standards, methods and tools planning This standard requires the identification of specific tools and methods to be used for development in general, not for each RISK CONTROL measure. |
| 14.7 Requirements specification For the PEMS and each of its subsystems (e.g. for a PESS) there shall be a documented requirement specification. | 5.2 Software requirements analysis This standard deals only with the software subsystems of a PEMS. |
| The requirement specification for a system or subsystem shall include and distinguish any ESSENTIAL PERFORMANCE and any RISK CONTROL measures implemented by that system or subsystem. | - Define and document software requirements from SYSTEM requirements. - Software requirements content - Include RISK CONTROL measures in software requirements This standard does not require that the requirements related to essential performance and RISK CONTROL measures be distinguished from other requirements, but it does require that all requirements be uniquely identified. |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 63 –

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.8 ARCHITECTURE For the PEMS and each of its subsystems, an ARCHITECTURE shall be specified that shall satisfy the requirements specification. | 5.3 Software ARCHITECTURAL design |
| Where appropriate, to reduce the RISK to an acceptable level, the architecture specification shall make use of: - COMPONENTS WITH HIGH-INTEGRITY CHARACTERISTICS; - fail-safe functions; - redundancy; - diversity; - partitioning of functionality; - defensive design, e.g. limits on potentially hazardous effects by restricting the available output power or by introducing means to limit the travel of actuators. | 5.3.5 Identify segregation necessary for RISK CONTROL Partitioning is the only technique identified, and it is only identified because there is a requirement to state how the integrity of the partitioning is assured. |
| The ARCHITECTURE specification shall take into consideration: - allocation of RISK CONTROL measures to subsystems and components of the PEMS; - failure modes of components and their effects; - common cause failures; - systemic failures; - test interval duration and diagnostic coverage; - maintainability; - protection from reasonably foreseeable misuse; - the IT-NETWORK specification, if applicable. | This is not included in this standard. |
| 14.9 Design and implementation Where appropriate, the design shall be decomposed into subsystems, each having both a design and test specification. | 5.4 Software detailed design 5.4.2 Develop detailed design for each SOFTWARE UNIT This standard does not require a test specification for detailed design. |
| Descriptive data regarding the design environment shall be included in the documentation. | 5.4.2 Develop detailed design for each SOFTWARE UNIT |
| 14.10 VERIFICATION VERIFICATION is required for all functions that implement BASIC SAFETY, ESSENTIAL PERFORMANCE or RISK CONTROL measures. | 5.1.6 Software VERIFICATION planning VERIFICATION is required for each ACTIVITY |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 64 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| A VERIFICATION plan shall be produced to show how these functions shall be verified. The plan shall include: - at which milestone(s) VERIFICATION is to be performed on each function; - the selection and documentation of VERIFICATION strategies, ACTIVITIES, techniques, and the appropriate level of independence of the personnel performing the VERIFICATION; - the selection and utilization of VERIFICATION tools; - coverage criteria for VERIFICATION. | 5.1.6 Software VERIFICATION planning Independence of personnel is not included in this standard. It is considered covered in ISO 13485. |
| The VERIFICATION shall be performed according to the VERIFICATION plan. The results of the VERIFICATION activities shall be documented. | VERIFICATION requirements are in most of the ACTIVITIES. |
| 14.11 PEMS VALIDATION A PEMS VALIDATION plan shall include the validation of BASIC SAFETY and ESSENTIAL PERFORMANCE. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| Methods used for PEMS VALIDATION shall be documented | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| The PEMS VALIDATION shall be performed according to the PEMS VALIDATION plan. The results of the PEMS VALIDATION activities shall be documented. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| The person having the overall responsibility for the PEMS VALIDATION shall be independent of the design team. The MANUFACTURER shall document the rationale for the level of independence. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| No member of a design team shall be responsible for the PEMS VALIDATION of their own design. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| All professional relationships of the members of the PEMS VALIDATION team with members of the design team shall be documented in the RISK MANAGEMENT FILE. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| A reference to the methods and results of the PEMS VALIDATION shall be included in the RISK MANAGEMENT FILE. | This standard does not cover software validation. PEMS validation is a SYSTEM level ACTIVITY and is outside the scope of this standard. |
| 14.12 Modification If any or all of a design results from a modification of an earlier design then either all of this clause applies as if it were a new design or the continued validity of any previous design documentation shall be assessed under a documented modification/change PROCEDURE. | 6 Software maintenance PROCESS This standard takes the approach that software maintenance should be planned and that implementation of modifications should use the software development PROCESS or an established software maintenance PROCESS. |
| When software is modified, the requirements in subclause 4.3, Clause 5, Clause 7, Clause 8 and Clause 9 of IEC 62304:2006 shall also apply to the modification. |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 65 –

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| 14.13 PEMS intended to be incorporated into an IT-NETWORK If the PEMS is intended to be incorporated into an IT-NETWORK that is not validated by the PEMS MANUFACTURER, the MANUFACTURER shall make available instructions for implementing such connection including the following - the purpose of the PEMS’S connection to an IT- NETWORK; - the required characteristics of the IT-NETWORK incorporating the PEMS; - the required configuration of the IT-NETWORK incorporating the PEMS; - the technical specifications of the network connection of the PEMS including security specifications; - the intended information flow between the PEMS, the IT-NETWORK and other devices on the IT-NETWORK, and the intended routing through the IT-NETWORK; and NOTE 1 This can include aspects of effectiveness and data and system security as related to BASIC SAFETY and ESSENTIAL PERFORMANCE (see also Clause H.6 and IEC 80001-1:2010). - list the HAZARDOUS SITUATIONS resulting from a failure of the IT-NETWORK to provide the characteristics required to meet the purpose of the PEMS connection to the IT-NETWORK. In the ACCOMPANYING DOCUMENTS, the MANUFACTURER shall instruct the RESPONSIBLE ORGANIZATION that: - connection of the PEMS to an IT-NETWORK that includes other equipment could result in previously unidentified RISKS to PATIENTS, OPERATORS or third parties; - the RESPONSIBLE ORGANIZATION should identify, analyze, evaluate and control these RISKS; | Requirements for incorporation into an IT-network are not included in this standard. |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 66 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

| PEMS requirements from IEC 60601-1:2005 | Requirements of IEC 62304 relating to the software subsystem of a PEMS |
| --- | --- |
| NOTE 3 IEC 80001-1:2010 provides guidance for the RESPONSIBLE ORGANIZATION to address these risks. - subsequent changes to the IT-NETWORK could introduce new RISKS and require additional analysis; and - changes to the IT-NETWORK include: - changes in the IT-NETWORK configuration; - connection of additional items to the IT- NETWORK; - disconnecting items from the IT- NETWORK; - update of equipment connected to the IT- NETWORK; - upgrade of equipment connected to the IT-NETWORK. |  |

#### Relationship to requirements in IEC 60601-1-4

IEC 60601-1-4 has been withdrawn.

### Relationship to IEC 61010-1

The scope of IEC 61010-1 [5] covers electrical test and measuring equipment, electrical control equipment and electrical laboratory equipment. Only part of the laboratory equipment is used in   a medical environment or as in vitro diagnostic equipment  (IVD).

Due to legal regulations or normative references, IVD equipment is allocated  to  MEDICAL  DEVICES without, however, falling within the scope of IEC 60601-1 [1]. This is attributable not  only to the fact that, strictly speaking, IVD instruments are not  MEDICAL  DEVICES  which  come into direct contact with patients, but also to the fact that such products are manufactured for many different applications in various laboratories. Use as an IVD instrument or  as  an accessory for an IVD instrument is then rare.

If laboratory equipment is used as IVD equipment, the measured results obtained must be EVALUATED in  accordance with medical criteria. The application of  ISO 14971 is required for  RISK MANAGEMENT. If such products also contain software that can lead to a HAZARDOUS  SITUATION, for example failure caused by the software which results in an unwanted change of medical data (measuring results), IEC 62304 must be taken into  account.

IEC 61010-1:2010 has a general requirement for risk assessment in Clause 17, which is more streamlined than the full risk management requirements of ISO 14971. Applying IEC 61010-1 Clause 17 alone does not meet the required criteria for risk management of IEC 62304, which     is based on full ISO 14971 risk management requirements. With this in mind, it is expected by this standard that when an IVD medical device has software-related risks, its risk management process is performed following ISO 14971 instead of only Clause 17 of IEC 61010-1. Compliance with Clause 17 of IEC 61010-1 will be achieved, as detailed in the Note to Clause    17 of IEC 61010-1:

NOTE  One RISK  assessment procedure is outlined in Annex J. Other RISK  assessment procedures are contained in ISO 14971, SEMI S10-1296, IEC 61508, ISO 14121-1, and ANSI B11.TR3. Other established procedures which implement similar steps can also be  used.

The flowchart in Figure C.3 shows the application of IEC 62304 with IEC 61010-1, Clause   17:

IEC 62304:2006	– 67 –

+AMD1:2015 CSV  IEC 2015

Intended purpose and use defined

Possible sources of HAZARD

identified

HAZARD related to the handling of medical data

Identify known and reasonably foreseeable

HAZARDS

Is the HAZARD covered by relevant safety standards?

Yes

Verify according to the relevant safety standard

No

Use

Yes

Does the device	No

ISO 14971

for RISK MANAGEMENT

provide medical relevant data

Does

the software have	No

any impact on the medical

data?

Select an applicable method for RISK CONTROL

based on safety standard

Additional requirements necessary to ensure that wrong data are detected prior to use of data for medical purposes.

Yes

Use of procedures required to verify the data?

Yes

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

No

Use IEC 62304

IEC  727/06

#### Figure C.3 – Application of IEC 62304 with IEC  61010-1

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 68 –	IEC 62304:2006

+AMD1:2015 CSV  IEC 2015

### Relationship to ISO/IEC 12207

This standard has been derived from the approach and concepts of ISO/IEC 12207 [9], which defines requirements for software life cycle PROCESSES in general, i.e. not restricted to MEDICAL DEVICES.

This standard differs from ISO/IEC 12207 mainly with respect to the following.  It:

- excludes SYSTEM aspects, such as SYSTEM requirements, SYSTEM ARCHITECTURE and validation;

- omits some PROCESSES seen as duplicating ACTIVITIES documented elsewhere for MEDICAL DEVICES;

- adds the (SAFETY) RISK MANAGEMENT PROCESS  and the software release  PROCESS;

- incorporates the documentation and the VERIFICATION supporting PROCESSES into the development and maintenance PROCESSES;

- merges the PROCESS  implementation and planning ACTIVITIES of each PROCESS  into a  single

ACTIVITY in the development and maintenance  PROCESSES;

- classifies the requirements with respect to SAFETY needs;  and

- does not explicitly classify PROCESSES as primary or supporting, nor group PROCESSES as ISO/IEC 12207 does.

Most  of  these changes  were  driven by the  desire to  tailor  the  standard to  the need of      the

MEDICAL DEVICE  sector by:

- focusing on SAFETY aspects and the MEDICAL DEVICE RISK MANAGEMENT standard ISO    14971;

- selecting the appropriate PROCESSES useful in a regulated  environment;

- taking into account that software development is embedded in a  quality system  (which covers some of the PROCESSES and requirements of ISO/IEC 12207);  and

- lowering the level of abstraction to make it easier to  use.

This standard is not contradictory to ISO/IEC 12207. ISO/IEC 12207 can be useful as an aide     in setting up a well structured SOFTWARE DEVELOPMENT LIFE CYCLE MODEL that includes the requirements of this standard.

Table C.5, which  was prepared  by  ISO/IEC  JTC1/SC7,  shows  the  relationship  between  IEC 62304 and ISO/IEC 12207.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 69 –

+AMD1:2015 CSV  IEC  2015

#### Table C.5 – Relationship to ISO/IEC  12207:2008

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
| 5 Software development PROCESS |  |  |  |
| 5.1 Software development planning |  |  |  |
|  | 5.1.1 Software development plan | 7.1.1 Software Implementation | 7.1.1.3.1 Software implementation strategy 7.1.1.3.1.1 7.1.1.3.1.3 7.1.1.3.1.4 6.3.1.3.2 Project planning 6.3.1.3.2.1 |
|  | 5.1.2 Keep software development plan updated | 6.3.2 Project Assesment and Control | 6.3.2.3.2 Project control 6.3.2.3.2.1 |
|  | 5.1.3 Software development plan reference to SYSTEM design and development | 6.4.3 System Architectural Design 6.4.5 System Integration 7.2.5 Software Validation Process | 6.4.3.3.1 Establishing architecture 6.4.3.3.1.1 6.4.5.3.1 Integration 6.4.5.3.1.1 7.2.5.3.1 Process implementation 7.2.5.3.1.4 |
|  | 5.1.4 Software development standards, methods and tools planning | 7.1.1 Software Implementation | 7.1.1.3.1 Software implementation strategy 7.1.1.3.1.3 |
|  | 5.1.5 Software integration and integration testing planning | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.1 |
|  | 5.1.6 Software VERIFICATION planning | 7.2.4 Software Verification - Software Construction - Software Integration - Software Qualification Testing | 7.2.4.3.1 Process implementation 7.2.4.3.1.4 7.2.4.3.1.5 7.1.5.3.1 Software constrution 7.1.5.3.1.5 7.1.6.3.1 Software integration 7.1.6.3.1.5 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.1.7 Software RISK MANAGEMENT planning | 6.3.4 Risk Management Process |  |
|  | 5.1.8 Documentation planning | 7.2.1 Software Documentation Management | 7.2.1.3.1 Process implementation 7.2.1.3.1.1 |
|  | 5.1.9 Software configuration management planning | 7.2.2 Software Configuration Management 7.2.8 Software Problem Resolution | 7.2.2.3.1 Process implementation 7.2.2.3.1.1 7.2.8.3.1 Process implementation 7.2.8.3.1.1 |
|  | 5.1.10 Supporting items to be controlled | 6.2.2 Infrastructure Management 6.2.2 Infrastructure Management | - Establishment of the infrastructure 6.2.2.3.2.1 - Maintenance of the infrastructure 6.2.2.3.3.1 |
|  | 5.1.11 Software CONFIGURATION ITEM control before VERIFICATION | 7.2.2 Software Configuration Management | 7.2.2.3.2 Configuration identification 7.2.2.3.2.1 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 70 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
| 5.2 Software requirements analysis |  |  |  |
|  | 5.2.1 Define and document software requirements from SYSTEM requirements | 6.4.3 System Architectural Design | 6.4.3.3.1 Establishing architecture 6.4.3.3.1.1 |
|  | 5.2.2 Software requirements content | 7.1.2 Software Requirements Analysis | 7.1.2.3.1 Software requirements analysis 7.1.2.3.1.1 |
|  | 5.2.3 Include RISK CONTROL measures in software requirements |  |  |
|  | 5.2.4 Re-EVALUATE MEDICAL DEVICE RISK ANALYSIS | None | None |
|  | 5.2.5 Update SYSTEM requirements | 7.1.2 Software Requirements Analysis | 7.1.2.3.1 Software requirements analysis 7.1.2.3.1.1 a) & b) |
|  | 5.2.6 Verify software requirements | 7.2.4 Software Verification | 7.2.4.3.2 Verification 7.2.4.3.2.1 |
| 5.3 Software ARCHITECTURAL design | 5.3.1 Transform software requirements into an ARCHITECTURE | 7.1.3 Software Architectural Design | 7.1.3.3.1 Software architectural design 7.1.3.3.1.1 |
|  | 5.3.2 Develop an ARCHITECTURE for the interfaces of SOFTWARE ITEMS |  | 7.1.3.3.1 Software architectural design 7.1.3.3.1.2 |
|  | 5.3.3 Specify functional and performance requirements of SOUP item | None | none |
|  | 5.3.4 Specify SYSTEM hardware and software required by SOUP item | None | none |
|  | 5.3.5 Identify segregation necessary for RISK CONTROL | None | none |
|  | 5.3.6 Verify software ARCHITECTURE | 7.1.3 Software Architectural Design | 7.1.3.3.1 Software architectural design 7.1.3.3.1.6 |
| 5.4 Software detailed design | 5.4.1 Refine SOFTWARE ARCHITECTURE into SOFTWARE UNITS | 7.1.4 Software Detailed Design | 7.1.4.3.1 Software detailed design 7.1.4.3.1.1 |
|  | 5.4.2 Develop detailed design for each SOFTWARE UNIT |  |  |
|  | 5.4.3 Develop detailed design for interfaces |  | 7.1.4.3.1 Software detailed design 7.1.4.3.1.2 |
|  | 5.4.4 Verify detailed design | 7.1.4 Software Detailed Design | 7.1.4.3.1 Software detailed design 7.1.4.3.1.7 |
| 5.5 SOFTWARE UNIT implementation and verification | 5.5.1 Implement each SOFTWARE UNIT | 7.1.5 Software Construction | 7.1.5.3.1 Software construction 7.1.5.3.1.1 |
|  | 5.5.2 Establish SOFTWARE UNIT VERIFICATION PROCESS | - Software Detailed Design - Software Construction | 7.1.4.3.1 Software detailed design 7.1.4.3.1.5 7.1.5.3.1 Software Construction 7.1.5.3.1.5 |
|  | 5.5.3 SOFTWARE UNIT acceptance criteria | 7.1.5 Software Construction | 7.1.5.3.1 Software construction 7.1.5.3.1.5 |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 71 –

+AMD1:2015 CSV  IEC  2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
|  | 5.5.4 Additional SOFTWARE UNIT acceptance criteria | 7.1.5 Software Construction 7.2.4 Software Verification | 7.1.5.3.1 Software construction 7.1.5.3.1.2 |
|  | 5.5.5 SOFTWARE UNIT VERIFICATION | 7.1.5 Software Construction | 7.1.5.3.1 Software construction 7.1.5.3.1.2 |
| 5.6 Software integration and integration testing | 5.6.1 Integrate SOFTWARE UNITS | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 |
|  | 5.6.2 Verify software integration | 7.1.6 Software Integration 6.4.5 System Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 6.4.5.3.1 Integration 6.4.5.3.1.2 |
|  | 5.6.3 Test integrated software | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.1 |
|  | 5.6.4 Integration testing content | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.6.5 Verify integration tests procedures | None | None |
|  | 5.6.6 Conduct regression tests | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 |
|  | 5.6.7 Integration test record contents | 7.1.6 Software Integration | 7.1.6.3.1 Software integration 7.1.6.3.1.2 |
|  | 5.6.8 Use software problem resolution PROCESS | 7.2.4 Software Verification | 7.2.4.3.1 Process implementation 7.2.4.3.1.6 |
| 5.7 SOFTWARE SYSTEM testing | 5.7.1 Establish tests for each software requirement | - Software Integration - Software Qualification Testing | 7.1.6.3.1 Software integration 7.1.6.3.1.4 7.1.7.3.1 Software qualification testing 7.1.7.3.1.1 |
|  | 5.7.2 Use software problem resolution PROCESS | 7.2.4 Software Verification | 7.2.4.3.1 Process implementation 7.2.4.3.1.6 |
|  | 5.7.3 Retest after changes | 7.2.8 Software Problem Resolution | 7.2.8.3.1 Process implementation 7.2.8.3.1.1 |
|  | 5.7.4 Verify SOFTWARE SYSTEM testing | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.7.5 SOFTWARE SYSTEM test record contents | 7.1.7 Software Qualification Testing | 7.1.7.3.1 Software qualification testing 7.1.7.3.1.1 |
| 5.8 Software release | 5.8.1 Ensure software VERIFICATION is complete | 6.4.9 Software Operation 7.2.2 Software Configuration Management | 6.4.9.3.2 Operation activation and check-out 6.4.9.3.2.1 6.4.9.3.2.2 7.2.2.3.6 Release management and delivery 7.2.2.3.6.1 |
|  | 5.8.2 Document known residual ANOMALIES | 7.2.2 Software Configuration Management 7.1.7 Software Qualification Testing | 7.2.2.3.5 Configuration evaluation 7.2.2.3.5.1 7.1.7.3.1 Software qualification testing 7.1.7.3.1.3 |
|  | 5.8.3 EVALUATE known residual ANOMALIES |  |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 72 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
|  | 5.8.4 Document released VERSIONS | 7.2.2 Software Configuration Management Process | 7.2.2.3.6 Release management and delivery 7.2.2.3.6.1 |
|  | 5.8.5 Document how released software was created |  |  |
|  | 5.8.6 Ensure ACTIVITIES and TASKS are complete |  |  |
|  | 5.8.7 Archive software |  |  |
|  | 5.8.8 Assure repeatability of software release |  |  |
| 6 Software maintenance PROCESS |  | 6.4.10 Software Maintenance Process |  |
| 6.1 Establish software maintenance plan |  | 6.4.10 Software Maintenance | None |
| 6.2 Problem and modification analysis | 6.2.1 Document and EVALUATE feedback | None | None |
|  | 6.2.1.1 Monitor feedback | 6.4.10 Software Maintenance | None |
|  | 6.2.1.2 Document and EVALUATE feedback |  |  |
|  | 6.2.1.3 EVALUATE PROBLEM REPORT’S effects on SAFETY | 6.4.10 Software Maintenance | None |
|  | 6.2.2 Use software problem resolution PROCESS | 6.4.10 Software Maintenance | None |
|  | 6.2.3 Analyse CHANGE REQUESTS | 6.4.10 Software Maintenance | None |
|  | 6.2.4 CHANGE REQUEST approval | 6.4.10 Software Maintenance | None |
|  | 6.2.5 Communicate to users and regulators | 6.4.10 Software Maintenance | None |
| 6.3 Modification implementation |  | None | None |
|  | 6.3.1 Use established PROCESS to implement modification | 6.4.10 Software Maintenance | None |
|  | 6.3.2 Re-release modified SOFTWARE SYSTEM | 7.2.2 Software Configuration Management | None |
| 7 Software RISK MANAGEMENT PROCESS |  | 6.3.4 Risk Management Process This is based on ISO/IEC 16085. While there is some commonality it does not address the specific requirements for medical device software development with regard to risk management |  |
| 8 Software configuration management PROCESS |  |  |  |
| 8.1 Configuration identification | 8.1.1 Establish means to identify CONFIGURATION ITEMS | 7.2.2 Software Configuration Management | None |
|  | 8.1.2 Identify SOUP | None | None |
|  | 8.1.3 Identify SYSTEM configuration documentation | 7.2.2 Software Configuration Management | None |
| 8.2 Change control | 8.2.1 Approve CHANGE REQUESTS | 7.2.2 Software Configuration Management | None |
|  | 8.2.2 Implement changes | 6.4.10 Software Maintenance | None |
|  | 8.2.3 Verify changes | 7.2.2 Software Configuration Management | None |
|  | 8.2.4 Provide means for TRACEABILITY of change |  |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 73 –

+AMD1:2015 CSV  IEC  2015

| ISO/IEC 62304 PROCESSES |  | ISO/IEC 12207:2008 |  |
| --- | --- | --- | --- |
| ACTIVITY | TASK | PROCESSES | ACTIVITY/TASK |
| 8.3 Configuration status accounting |  | 7.2.2 Software Configuration Management | Non e |
| 9 Software problem resolution PROCESS |  |  |  |
| 9.1 Prepare PROBLEM REPORTS |  | 7.2.8 Software Problem Resolution | None |
| 9.2 Investigate the problem |  | 7.2.8 Software Problem Resolution | None |
| 9.3 Advise relevant parties |  | 7.2.8 Software Problem Resolution | None |
| 9.4 Use change control process |  | 7.2.2 Software Configuration Management 6.4.10 Software Maintenance | None |
| 9.5 Maintain records |  | 7.2.8 Software Problem Resolution | None |
| 9.6 Analyse problems for trends |  | 7.2.8 Software Problem Resolution | None |
| 9.7 Verify software problem resolution |  | 7.2.8 Software Problem Resolution | None |
| 9.8 Test documentation contents |  | All testing TASKS in ISO 12207 require documentation | None |

### C.7     Relationship to IEC 61508

The question has been raised whether this standard, being concerned with the design of SAFETY-critical software, should follow the principles of IEC 61508. The approach to safety in IEC 62304 is fundamentally different than the one in IEC 61508.  IEC 62304  takes  into account that the effectiveness of medical devices justifies residual risks related to  their use. The following explains the stance of this  standard.

IEC 61508 addresses 3 main  issues:

- RISK MANAGEMENT life cycle and life cycle  PROCESSES;

- definition of Safety Integrity Levels;

- recommendation of techniques, tools and methods for software development and levels of independence of personnel responsible for performing different  TASKS.

Issue 1) is covered in this standard by a normative reference to  ISO  14971 (the MEDICAL DEVICE sector standard for RISK MANAGEMENT). The effect of this reference is to adopt ISO 14971’s  approach  to  RISK  MANAGEMENT   as  an  integral  part  of  the  software  PROCESS    for

MEDICAL DEVICE SOFTWARE.

For issue 2), this standard takes a simpler approach than IEC 61508. The latter classifies software into 4 “Safety Integrity Levels” defined in terms of  reliability  objectives.  The reliability objectives are identified after RISK ANALYSIS, which quantifies both the severity and the probability of HARM caused by a failure of the   software.

This standard simplifies issue 2) by defining the classification into 3 software safety classes based on the RISK caused by a failure.  After classification, different PROCESSES  are required  for different software safety classes: the intention is to further reduce the probability (and/or  the severity) of failure of the  software.

Issue 3) is not addressed by this standard. Readers of the standard are encouraged to use    IEC 61508 as a source for good software methods, techniques and tools, while recognising

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 74 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

that other approaches, both present and future, can provide equally good results.  This standard makes no recommendation concerning independence of people responsible for one software ACTIVITY (for example VERIFICATION) from those responsible for another (for example design). In particular, this standard makes no requirement for an  independent  safety assessor, since this is a matter for ISO  14971.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 75 –

+AMD1:2015 CSV  IEC  2015

# Annex D

## (informative)

Implementation

### Introduction

This annex gives an overview of how this standard can be implemented into MANUFACTURERS’ PROCESSES. It also considers that other standards like ISO 13485 [8] require adequate and comparable PROCESSES.

### Quality management system

For MANUFACTURERS of MEDICAL DEVICES, including MEDICAL DEVICE SOFTWARE in the context of this standard, the establishment of a quality management system (QMS) is  required in  4.1.  This standard does not require that the QMS necessarily has to be   certified.

  - EVALUATE quality management PROCESSES

It is recommended to EVALUATE how well the established and documented PROCESSES of the QMS already cover the PROCESSES of the software life cycle, by means of audits, inspections, or analyses under the responsibility of the MANUFACTURER. Any identified gaps can be accommodated by extending the QM PROCESSES, or can be separately described. If the MANUFACTURER already has PROCESS descriptions available which regulate the development, VERIFICATION and validation of software, then these should also be EVALUATED  to determine how well they agree with this  standard.

  - Integrating requirements of this standard into the MANUFACTURER’S quality management PROCESSES

This standard can be implemented by adapting or extending the PROCESSES  already installed   in the QMS system, or integrating new PROCESSES. This standard does not specify how this is   to be done; the MANUFACTURER is free to do this in any suitable   way.

The MANUFACTURER is responsible for ensuring that the PROCESSES described in this standard are suitably put into action when the MEDICAL DEVICE SOFTWARE is developed by Original Equipment Manufacturers (OEM) or sub-contractors not having their own documented   QMS.

  - Checklist for small MANUFACTURERS without a certified  QMS

The MANUFACTURER should determine the highest software safety classification (A, B or C) of the software. Table D.1 lists  all  ACTIVITIES  described  in  this  standard.  The  reference  to ISO 13485 should help to define the place in the QMS. Based on the required software safety class,  the   MANUFACTURER   should  assess  each  required   ACTIVITY   against  the  existing

PROCESSES.  If  the  requirement  is  already  covered,  a  reference  to  the  relevant    PROCESS

descriptions should be given.

If there is discrepancy, an action is needed to improve the   PROCESS.

The list can also be used for an EVALUATION of the PROCESSES after the action has been performed.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 76 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

#### Table D.1 – Checklist for small companies without a certified   QMS

| ACTIVITY | Related clause of ISO 13485:2003 | Covered by existing procedure? | If yes: Reference | Actions to be taken |
| --- | --- | --- | --- | --- |
| 5.1 Software development planning | 7.3.1 Design and development planning | Yes/No |  |  |
| 5.2 Software requirements analysis | 7.3.2 Design and development inputs | Yes/No |  |  |
| 5.3 Software ARCHITECTURAL design |  | Yes/No |  |  |
| 5.4 Software detailed design |  | Yes/No |  |  |
| 5.5 SOFTWARE UNIT implementation and verification |  | Yes/No |  |  |
| 5.6 Software integration and integration testing |  | Yes/No |  |  |
| 5.7 SOFTWARE SYSTEM testing | - Design and development outputs - Design and development review | Yes/No |  |  |
| 5.8 Software release | - Design and development verification - Design and development validation | Yes/No |  |  |
| 6.1 Establish software maintenance plan | 7.3.7 Control of design and development changes | Yes/No |  |  |
| 6.2 Problem and modification analysis |  | Yes/No |  |  |
| 6.3 Modification implementation | - Design and development verification - Design and development validation | Yes/No |  |  |
| 7.1 Analysis of software contributing to hazardous situations |  | Yes/No |  |  |
| 7.2 RISK CONTROL measures |  | Yes/No |  |  |
| 7.3 VERIFICATION of RISK CONTROL measures |  | Yes/No |  |  |
| 7.4 RISK MANAGEMENT of software changes |  | Yes/No |  |  |
| 8.1 Configuration identification | 7.5.3 Identification and traceability | Yes/No |  |  |
| 8.2 Change control | 7.5.3 Identification and traceability | Yes/No |  |  |
| 8.3 Configuration status accounting |  | Yes/No |  |  |
| 9 Software problem resolution PROCESS |  | Yes/No |  |  |

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

IEC 62304:2006	– 77 –

+AMD1:2015 CSV  IEC  2015

## Bibliography

[1]		IEC 60601-1:2005,  Medical  electrical  equipment  –  Part  1:  General  requirements  for basic safety and essential  performance

IEC 60601-1:2005/AMD1:2012

[2]	IEC 60601-1-4:1996, Medical electrical equipment –  Part 1:  General    requirements for safety – 4.Collateral standard: Programmable electrical medical systems  (withdrawn)

IEC 60601-1-4:1996/AMD1:1999

[3]	IEC 60601-1-6, Medical electrical equipment – Part 1-6: General requirements   for basic safety and essential performance – Collateral standard:  Usability

[4]	IEC 61508-3, Functional safety of  electrical/electronic/programmable   electronic safety- related systems – Part 3: Software  requirements

[5]		IEC 61010-1:2010,  Safety  requirements  for   electrical  equipment   for  measurement, control, and laboratory use – Part 1: General  requirements

[6]	ISO 9000:2005, Quality management systems – Fundamentals and  vocabulary

[7]	ISO 9001:2008, Quality management systems –  Requirements

[8]		ISO  13485:2003, Medical devices –  Quality management systems –    Requirements for regulatory purposes

[9]		ISO/IEC  12207:2008,   Systems  and   software   engineering   –   Software   life   cycle processes

[10]	ISO/IEC  14764:1999,   Software   Engineering   –   Software   Life  Cycle  Processes –

Maintenance

[11]	ISO/IEC  15504-5:2012,  Information  technology  –  Process  assessment  –  Part  5:  An exemplar software life cycle process assessment  model

[12]		ISO/IEC 25010:2011, Systems and software engineering – System and   software Quality Requirements and Evaluation (SQuaRE) – System and software quality   models

[13]	ISO/IEC 33001:—2),  Information  technology  –  Process  assessment  –   Concepts  and terminology

[14]	ISO/IEC 33004:—2), Information technology –  Process assessment –   Requirements for process reference, process assessment and maturity  models

[15]	ISO/IEC  90003:2014,   Software   engineering   –   Guidelines   for   the   application  of ISO 9001:2008 to computer software

[16]	ISO/IEC Guide 51:2014, Safety aspects – Guidelines for their inclusion in   standards

[17]	IEEE 610.12:1990, IEEE standard glossary of software engineering  terminology

[18]	IEEE 1044:2009, IEEE standard classification for software  anomalies

- To be published.

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

– 78 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

[19] U.S. Department Of Health and Human Services, Food and Drug Administration, Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices, May 11, 2005,

<://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/u cm089543.htm>

[20] U.S. Department Of Health and Human Services, Food  and  Drug  Administration, General Principles of Software Validation; Final Guidance for Industry and FDA Staff, January 11, 2002,

<http://www.fda.gov/downloads/RegulatoryInformation/Guidances/ucm126955.pdf>

[21] IEC 62366-1:2015, Medical devices – Part 1: Application of usability engineering to medical devices

[22]	IEC 82304-1:—3),   Healthcare Software Systems – Part 1: General  requirements

- In preparation.

IEC 62304:2006	– 79 –

+AMD1:2015 CSV  IEC  2015

# Index of defined terms

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

ACTIVITY, 15, 17, 23, 25, 27, 31, 33, 43, 59,    65,

67, 69, 73, 79, 81, 83, 87, 89, 95, 113, 133,

145

Change control, 101

Change request, 61

Completion of, 49

Configuration identification, 101

Configuration management, 35

Configuration status accounting, 101 Definition, 19

Deliverable, 19

Design and maintenance, 11 Hazard identification, 11

Maintenance, 51

Mapping, 15

Modification implementation, 97

Planning, 83, 85

Problem and modification analysis, 95 Problem resolution, 31, 53, 103

Required, 15, 147

Requirements, 17

Requirements analysis, 39

Risk analysis, 55

Risk management, 33, 47, 59, 79, 81, 99 Software architectural design, 87

Software detailed design, 89 Software development, 11

Software integration, 93

Software integration and integration testing, 91

Software maintenance, 95

Software release, 95

Software requirements analysis, 85 Software system testing, 93

SOFTWARE UNIT implementation and verification, 89

Testing, 45, 47

Verification, 33

ANOMALY, 45, 47, 49, 55, 65,  93

Definition, 19

ARCHITECTURE, 39, 41, 73, 75, 79, 81, 83, 85,

87, 89, 99, 113, 133

Definition, 19

CHANGE REQUEST, 53, 61, 63, 65, 97,  101

Definition, 19

CONFIGURATION  ITEM,  27,  35,  49,  59,  61,   97,

101

Definition, 19

SOUP, 31, 59

DELIVERABLE, 25, 31, 33

Definition, 19

EVALUATION, 41, 45, 49, 51, 53, 55, 57, 87,    89,

93, 95, 99, 147, 149

Re-, 39

HARM, 21, 23, 73, 81,  145

Definition, 21

HAZARD, 11, 23, 29, 57, 67, 69, 79, 83, 93, 97,

99, 129

Definition, 21

Unforeseen, 87

MANUFACTURER, 15, 21, 23, 25, 27, 29, 31,    33,

35, 37, 39, 41, 43, 45, 47, 49, 51, 53, 55, 57,

59, 61, 63, 65, 75, 77, 79, 81, 83, 85, 87, 89,

91, 93, 95, 97, 101, 103, 107,  147

Definition, 21

MEDICAL DEVICE, 11, 17, 21, 27, 35, 39, 41,    55,

69,  75,  77,  79,  85,  87,  91,  93,  95,  97, 99,

105, 129, 133, 145, 147

Definition, 21

MEDICAL  DEVICE  SOFTWARE, 11,  13,  17, 27, 35,

37, 39, 51, 67, 73, 75, 77, 79, 83, 85, 91, 93,

95, 97, 101, 105, 145, 147

Change, 59

Definition, 21

PROBLEM REPORT, 51, 53, 61, 63, 65, 95,  97

Classification, 61 Definition, 21

PROCESS,  13,  15,  17,  23,  25,  27,  31,  67, 69,

73, 75, 79, 81, 85, 87, 89, 97, 101, 103,    113,

133, 145, 147

Acceptance, 61

Change control, 61, 63 Classification, 133

Configuration management, 51, 89, 113

Decision-making, 77

Definition, 23

Development, 27, 81, 95, 113

Existing, 31

Improvement, 149

Life cycle, 11, 133, 143

Maintenance, 51, 53, 113

Mapping, 15

Modification, 97

Omission of, 81

Output, 75

Physiological, 21

Problem resolution, 35, 45, 47, 51, 53,  63,

97, 101, 103, 113

Quality management, 147

Required, 15, 147

Requirements, 17, 29

Risk analysis, 73

Risk management, 11, 23, 29, 33, 51, 63,  79,

81, 85, 89, 99, 109, 113, 129,  133

Software, 79, 145

Software development, 11, 27, 31, 53,  73

Software maintenance, 11, 95, 97

Software release, 133

System requirements, 87

Verification, 27

REGRESSION TESTING, 45, 65, 93

Definition, 23

RISK, 23, 67, 75, 79, 81, 83, 85, 91, 97,   99

Definition, 23

Non-serious injury, 29

Reasonably foreseeable, 79

Risk control, 23

Serious injury, 29

SOUP, 33

Unacceptable, 11, 25, 49

– 80 –	IEC 62304:2006

+AMD1:2015 CSV  IEC  2015

SOFTWARE  SYSTEM, 21, 25, 29, 31, 33, 37,     43,

53, 59, 61, 69, 73, 77, 79, 81, 83, 85, 89, 93,

95, 111

Definition, 25

RISK ANALYSIS, 39, 55, 67, 73, 79, 87, 99,  145

Definition, 23 RISK CONTROL

Activity, 11

Definition, 23

Hardware measure, 29

Measure, 29, 31, 37, 43, 45, 55, 57, 59,  79,

81, 85, 87, 89, 93, 97,  99

Requirements, 39, 41, 57, 99

Segregation, 41

RISK  MANAGEMENT, 11,  23,  29,  33,  47, 51, 53,

59,  63,  67,  75,  77,  79,  81,  85,  87,  89, 99,

109, 113, 129, 133, 145

Definition, 23

Medical device, 75

Report, 57

RISK  MANAGEMENT  FILE, 17, 29, 55, 57, 63,   87,

89, 97

Definition, 23

SAFETY, 11, 51, 63, 69, 77, 81, 89, 91, 93, 95,

97, 103, 133, 143

Definition, 25

SECURITY, 63

Definition, 25

Requirements, 37

SERIOUS INJURY, 29, 83

Definition, 25

Non-, 29, 83

SOFTWARE DEVELOPMENT LIFE CYCLE MODEL, 31, 73, 133

Definition, 25

SOFTWARE  ITEM, 25, 27, 29, 31, 33, 39, 41,    43,

53, 55, 57, 61, 65, 67, 69, 75, 77, 79, 81, 83,

87, 89, 91, 93, 97, 101,  111

Changed, 53

Definition, 25

Integration, 43, 45

Partitioning, 81

Performance, 45

Segregation, 41

SOUP, 27, 33, 39

Software Of Unknown Provenance See SOUP, 27

SOFTWARE PRODUCT, 19, 21, 23, 25, 27, 31,   49,

51, 53, 59, 61, 65, 73, 77, 85, 89, 91,   97

Definition, 25

Released, 51, 53

Integration, 43

Requirements, 35

Testing, 45, 47

SOFTWARE UNIT, 25, 41, 43, 73, 77, 89,  91

Definition, 27

Integration, 43

Verification, 43

SOFTWARE UNIT Verification, 41

SOUP, 33, 35, 39, 41, 51, 55, 59, 75,  85

Change, 59

Configuration item, 31

Definition, 27

Designator, 59

Software item, 33

SYSTEM, 11, 19, 21, 23, 25, 31, 37, 39, 65, 73,

75, 79, 83, 85, 87, 101,  133

Configuration, 61

Definition, 27

Development plan, 31

Existing, 51

Released, 53

Requirements, 33, 35, 39, 41

TASK, 15, 17, 19, 23, 25, 29, 31, 73, 83, 93,   95,

97, 143

Completion of, 49

Configuration management, 35

Definition, 27

Deliverable, 19

Design and maintenance, 11 Maintenance, 51

Mapping, 15

Required, 15

Requirements, 17

Risk management, 33

Verification, 33

TRACEABILITY, 31, 57, 85, 87

Definition, 27

Verification, 25, 33, 35, 41, 43, 47, 49, 57, 61,

63, 69, 73, 75, 87, 91, 93, 97, 101, 113,    133,

145, 147

Definition, 27

VERSION, 49, 55, 59, 65, 95,  101

Definition, 27

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

__________

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.

INTERNATIONAL

ELECTROTECHNICAL COMMISSION

3, rue de Varembé PO Box 131

CH-1211 Geneva 20

Switzerland

Tel:  + 41 22 919 02 11

Fax: + 41 22 919 03 00

info@iec.ch www.iec.ch

Copyrighted material licensed to Laurence Sampson on 2016-10-27 for licensee's use only.  No further reproduction or networking is permitted.  Distributed by Thomson Reuters (Scientific) LLC, www.techstreet.com.